## Slide 1: Astronomie et astrophysique pour physiciens CUSO 2012 Instruments and observational techniques – Spectroscopy
F. Pepe
Observatoire de l’Université Genève
F. Courbin and P. Jablonka, EPFL

## Slide 2: Course outline
PART 1 - Principles of spectroscopy
  Fundamental parameters
  Overview of spectrometry methods
PART 2 – ‘Modern’ spectrographs
  ‘Simple’ spectroimager
    FORS
  Echelle spectrographs
    UVES
    CRIRES
PART 3 - Spectroscopy on the VLTs
  Multi-Object spectrographs (MOS) and Intergral-Field Units (IFU) and spectro-imagers
    Giraffe
    VIMOS
    Sinfoni
  Future instruments
    X-shooter
    MUSE
    ESPRESSO

## Slide 3: The observables
If we integrate the surface brightness over a given source or sky aperture, we get the spectral flux density Fn or Fl at a given light frequency or wavelength:

## Slide 4: Filter spectrometer
Filter
Telescope
Focal plane
Objective
Detector (single or array)

## Slide 5: Filter spectrometer
Detector records Il for a given filter with transmittance tc, central wavelength lc, and band width Dl
tc, Dl and lc need to be calibrated on standard sources
Appropriate for broad-band spectra
Only one channel per measurement (unless dichroics are used)

## Slide 6: Filter spectrometer
Filter 1
Telescope
Objective
Detector (single or array)
Filter 2
Dichroic

## Slide 7: Fabry-Pérot spectrometer
Telescope
Focal plane
Objective
Detector
Collimator
Fabry-Pérot etalon
Slit

## Slide 8: Fabry-Pérot spectrometer
Similar to filter spectrometer, but spacing can be made tunable
Detector records I(l) as a function of the transmitted wavelength ml=2l, where m is an integer and enumerates the transmitted order.
Only one spectral channel per measurement
Transmittance and wavelength must be calibrated.
Allows high spectral resolution, if the finesse F or the order m is high. In the latter case, a pre-filtering is required to select only one wavelength (order-selection).

## Slide 9: Fabry-Pérot spectrometer
Order-separating filter (e.g. with second FP in 1st order)

## Slide 10: Fabry-Pérot spectrometer
Telescope
Focal plane
Objective
Detector
Collimator
Fabry-Pérot etalon
Slit
Real Fabry-Pérot:
D
f
D

## Slide 11: Example of apearture finesse
Transmitted wavelength is unique only for given angle T
If slit is too wide, the aperture (angle cone) is enlarged and the range of transmitted wavelength increased.
The ‘contrast is the reduced, thus the finesse and the spectral resolution

## Slide 12: General spectrograph layout
Telescope
Focal plane
Objective
Detector
Collimator
Disperser
Slit
Single detector -> monochromator (may be used with movable part to scan over wavelengths
Array detector -> spectrograph with N wavelength channels (N = number of detectors or pixels)

## Slide 13: Dispersers
The disperser separates the wavelengths in angular direction. To avoid angular mixing, the beam is collimated. The disperser is characterized by its angular dispersion:
where b is the deviation angle from the un-dispersed direction

## Slide 14: The prism

## Slide 15: Prism characteristics
Prism example: BK7 (normal glass), t = 50 mm, D = 100 mm
-> High transmittance
-> When used at minimum deviation, no anamorphosis (beam dia. compression or enlargement)
-> Produces ‘low’ dispersion

## Slide 16: Prism characteristics
-> Depends mainly on glass material (internal transmittance)
-> Anti-reflection coatings are needed to avoid reflection losses, especially for large apex angles (and large a). The coating must be optimized for the glass and the used angles.
-> Efficiency can be as high as 99%
-> The dispersion increases towards the blue wavelengths. For Crown glasses (contain Potassium) the ratio of the dispersion between blue and red is lower than for Flint glasses (contain lead, titanium dioxide or zirconium dioxide) .

## Slide 17: Grating spectrograph
Camera
Collimator
Focal ratios defined as
Fi = fi / Di

## Slide 18: The diffraction grating
Generic grating equation from the condition of positive interference between various ‘grooves’:

## Slide 19: Grating characteristics
Typical grating example: m = 1, r = 1000 gr/mm, sina + sinb = 1, cosb=1/2
-> Several orders result for a given wavelength
-> m = 0 for a grating which acts like a mirror -> no dispersion!
-> Orders overlap spatially -> must be filtered or use at m=1
Dispersion typically much higher than for prisms!

## Slide 20: Grism efficiency
Maximum efficiency obtained when specular groove reflection is matches (Blaze condition):

## Slide 21: Order overlaps
0
lU
2lL
2lU
lL
Wavelength in first order marking position on detector in dispersion direction (if dispersion ~linear)
Effective passband
in 1st order
1st order
blaze profile
Passband
in zero
order
m=1
m=2
m=0
2nd order
blaze profile
Intensity
Passband
in 2nd order
First and second
orders overlap!
Don't forget higher orders!
0
lL
lU
1st order
(2nd order)
lC
Zero order
matters for MOS

## Slide 22: Resolving power and resolution
Camera
Collimator
Focal ratios defined as
Fi = fi / Di

## Slide 23: Resolving power and resolution
The objective translates angles into positions on the detector. Each position (pixel) of the detector ‘sees’ a given angle of the parallel (collimated) beam
The collimated beam is never perfectly parallel, because either of the limited diameter of the beam, which produces diffraction df = 1.22 l/D1, or because of the finite slit, which produces and angular divergence dQ = s/f1.
The angular divergence is translated into a distance dl = f2 dQ or dl = f2 df on the CCD. This means that over this distances the wavelengths are mixed (cannot be separated angularly.

## Slide 24: Resolving power and resolution
Resolving power is the maximum spectral resolution which can be reached if the slit s = 0 and the angular divergence is limited by diffraction arising from the limited beam diameter. For a given Dispersion D we get the resolving power:
Spectral resolution is the effective spectral resolution which is finally reached when assuming a finite slit s. For a given Dispersion D we get the spectral resolution:

## Slide 25: Conservation of the ‘étendue’
The étendue is defined as E=A x O, where A is the area of the beam at a given optical surface and O is the solid angle under which the beam passes through the surface.
When following the optical path of the beam through an optical system, E is constant, in particular, it cannot be reduced
For a telescope, E is the product of the primary mirror surface and the two-dimensional field (in sterad) transmitted by the optical system. Normally, the transmitted field is defines a slit width. When entering spectrograph, the slit x beam aperture at the slit is equal to the etendue E of the telescope. This implies that at fixed spectral resolution, the slit width and the beam diameter cannot be chosen independently, since dQ depends on both.

## Slide 26: Other dispersers
Grisms
VPHG
Echelle grating

## Slide 27: Grisms
Transmission grating attached to prism
Allows in-line optical train:
  simpler to engineer
  quasi-Littrow configuration - no variable anamorphism
Inefficient for r > 600/mm due to groove shadowing and other effects

## Slide 28: Grism equations
Modified grating equation:
Undeviated condition:
  n'= 1, b = -a = f
Blaze condition: q = 0  lB = lU
Resolving power
(same procedure as for grating)
q = phase difference from centre of one ruling to its edge

## Slide 29: Volume Phase Holographic gratings
So far we have considered surface relief gratings
An alternative is VPH in which refractive index varies harmonically throughout the body of the grating:
Don't confuse with 'holographic' gratings (SR)
Advantages:
  Higher peak efficiency than SR
  Possibility of very large size with high r
  Blaze condition can be altered (tuned)
  Encapsulation in flat glass makes more robust
Disadvantages
  Tuning of blaze requires bendable spectrograph!
  Issues of wavefront errors and cryogenic use

## Slide 30: VPH configurations
Fringes = planes of constant n
Body of grating made from Dichromated Gelatine (DCG) which permanently adopts fringe pattern generated holographically
Fringe orientation allows operation in transmission or reflection

## Slide 31: VPH equations
Modified grating equation:
Blaze condition:
= Bragg diffraction
Resolving power:
Tune blaze condition by tilting grating (a)
Collimator-camera angle must also change by 2a  mechanical complexity

## Slide 32: VPH efficiency
Kogelnik's analysis when:
Bragg condition when:
Bragg envelopes (efficiency FWHM):
  in wavelength:
  in angle:
Broad blaze requires
  thin DCG
  large index amplitude
Superblaze
Barden et al. PASP 112, 809 (2000)

## Slide 33: Resolving power and resolution
The objective translates angles into positions on the detector. Each position (pixel) of the detector ‘sees’ a given angle of the parallel (collimated) beam
The collimated beam is never perfectly parallel, because either of the limited diameter of the beam, which produces diffraction df = 1.22 l/D1, or because of the finite slit, which produces and angular divergence dQ = s/f1.
The angular divergence is translated into a distance dl = f2 dQ or dl = f2 df on the CCD. This means that over this distances the wavelengths are mixed (cannot be separated angularly.

## Slide 34: Resolving power and resolution
Resolving power is the maximum spectral resolution which can be reached if the slit s = 0 and the angular divergence is limited by diffraction arising from the limited beam diameter. For a given Dispersion D we get the resolving power:
Spectral resolution is the effective spectral resolution which is finally reached when assuming a finite slit s. For a given Dispersion D we get the spectral resolution:

## Slide 35: Example of spectrographs
Basic Parameters
Telescope diameter: DT
Source/seeing/slit: ssky
Collimated beam of the spectrograph: D1
Other parameters:
Telescope focal length: fT
Telescope F-number (focal ratio): FT = F = fT/DT
Physical slit/fiber width: s = fT x ssky
Collimator focal length: f1
Objective focal length: f2

## Slide 36: Grating spectrograph
Camera
Collimator
Focal ratios defined as
Fi = fi / Di

## Slide 37: Example of simple spectrographs
Spectral resolution:
where

## Slide 38: Example of ‘simple’ spectrographs
Coralie@Euler:
Telescope diameter: DT = 1.2 m
Source/seeing/slit: ssky = 1 arcsec
Collimated beam of the spectrograph: D1 = 75 mm

## Slide 39: Example of ‘simple’ spectrographs
With prism: BK7 (normal glass), t = 50 mm
With grism: m = 1, r = 150 gr/mm, cosb=1

## Slide 40: FORS@VLT

## Slide 41: FORS@VLT

## Slide 42: FORS@VLT

## Slide 43: FORS@VLT

## Slide 44: FORS: Filter mode

## Slide 45: FORS: Grism mode
Grisms from 150 gr/mm to 1400 gr/mm
Spectroscopic modes
  ‘slitless’ MOS mode (R is given by seeing)
  Mask with up to 9 ‘long slits’ of 0.3” x 6.8’
  Up to 19 movable ‘slitlets’ of 0.3” x 22.5”
  MOS-MXU: Laser-cut masks (any format)
Spectral resolution depends on grism dispersion and slit width in dispersion direction.

## Slide 46: FORS: Grism mode
Movable slilets (localized sourcess in ‘crowded fields)
Long slit (extended sources, 1-D image + spectrum)

## Slide 47: FORS: MOS mode

## Slide 48: FORS@VLT

## Slide 49: FORS@VLT
This image shows the first spectrum obtained of the comet 1995 Q1. The spectral coverage extends from about 3650 Angstrom (left) to 4100 Angstrom (right) in the violet region. The spectrograph slit was oriented along the main tail at position angle ~ 145 degrees. The total slit length was 5.6 arcminutes. The spectrum is typical for a comet at Comet Bradfield's current distance from the Sun (0.5 A.U., or about 75 million kilometres).

## Slide 50: FORS: Spectral resolution
Telescope diameter: DT = 8.2 m
Source/seeing/slit: ssky = 0.3 arcsec
Collimated beam : D1 = 125 mm
With grism: m = 1, r = 1400 gr/mm, cosb=1

## Slide 51: Example of ‘simple’ spectrographs
With single prisms resolution remains small
Possibility to increase dispersion by adding up several prisms
With gratings or grisms, the spectral resolution is up to several 1000 but still modest.
Other ‘tricks’ required

## Slide 52: Echelle grating
Increase dispersion of the ordinary greating by increasing the difference between entrance and exit angle
The use in Littrow condition will make the mounting symmetric and maximize a + b, since a is set equal to b: a = b
Use in blaze condition for maximum efficiency: Groove angle g = b

## Slide 53: Echelle grating
From the grating equation we derive the dispersion law of an echelle grating:
Example: R = 4 grating:

## Slide 54: Characteristics of Echelle grating
High dispersion -> high resolution
Spectrum becomes VERY long and efficiency drops because of blaze function
Alternative: use at high order, e.g. m = 100 (r = 30 gr/mm)
-> Orders overlap spacially and must be separated (filtering, pre-dispersion or cross-dispersion)
Free spectral range:
Numbers of orders for full ‘octave’:

## Slide 55: Multiple orders
Many orders to cover desired ll: Free spectral range
Fl= l/m
Orders lie on top of each other:
l(m) = l(n) (n/m)
Solution:
  use narrow passband filter to isolate one order at a time
  cross-disperse to fill detector with many orders at once
Cross dispersion may use prisms or low dispersion grating
Blaze condition fulfilled

## Slide 56
Monochromatic image of the slit
Main dispersion (echelle grating) ->
Cross dispersion (prism, grism, grating) ->

## Slide 57: UVES@VLT

## Slide 58: UVES@VLT

## Slide 59: UVES@VLT

## Slide 60: Coralie@Euler

## Slide 61: UVES@VLT

## Slide 63: UVES: Spectral format

## Slide 64: UVES: Order separation

## Slide 65: UVES: Spectral resolution
Telescope diameter: DT = 8.2 m
Source/seeing/slit: ssky = 0.3 arcsec
Collimated beam of the spectrograph: D1 = 200 mm
Echelle grating: R4 -> tanb = 4, D1 = 200 mm

## Slide 66: UVES: Sampling and resolution

## Slide 67: UVES: Image slicer ‘trick’
Image slicer
Long slit
Increase of efficiency at give R, or vice versa

## Slide 68: UVES params

## Slide 69: UVES: Efficiency

## Slide 72: Flames@VLT

## Slide 73: Flames – various links
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315
Giraffe
UVES

## Slide 74: Flames – UVES link
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 75: Flames – UVES link
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 76: Flames – Giraffe

## Slide 77: Flames – MEDUSA
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 78: Flames – MEDUSA
OzPoz fiber positioner

## Slide 79: Flames – MEDUSA
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 80: Flames – IFU
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 81: Flames – IFU

## Slide 82: Flames – ARGUS
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 83: Flames – ARGUS
Mode Number of Buttons Fibers per Buttons Sky Buttons Total fibers
UVES 8 1 - 8
Medusa 132 1 - 132
IFU 15 20 15 315
ARGUS 1 14x22 (-8) 15 315

## Slide 84: Flames – ARGUS

## Slide 85: Flames – Summary

## Slide 86: Flames – Summary
MOS
IFU
Single field
Field
Slit
Spectrum

## Slide 87: Future VLT spectrographs
MUSE:
AO
IFU with 1’ x 1’ (0,2” sampling) or 7.5” x 7.5” (0,025” sampling)
Array of 24 fields organized in 24 spectrographs
Resolution 2000 – 4000 @ NIR
Available in 2011

## Slide 88: Challenges for ELT spectrographs
Huge telescope -> huge instrumentations (conservation of étendu)
Or, massive adaptive optics (very complex and demanding with increasing telescope size, difficulty increases with D and with ~1/l
Location of the instruments (trade-off between stability and efficiency)
Costs