Centro de Astrofísica da Universidade do Porto

Universidade de Lisboa, SIM/IDL & LOLS

INAF, Osservatorio Astronomico di Trieste

INAF, Osservatorio Astronomico di Brera

Observatory of the University of Geneva

Physics Institute, University of Bern

Instituto de Astrofísica de Canarias

European Southern Observatory

ESPRESSO

# Fabry-Perot Calibrator

# Final Design Description and Performances Analysis

# VLT-TRE-ESP-13520-xxx, Issue 0

# March 14th, 2013

Name				Date				Signature

Prepared	F. Pepe	14/03/2013

Name				Date				Signature

Approved	F. Zerbi

Name				Date				Signature

Released	F. Pepe

# Change Record

| Issue/Rev. | Date | Section/Page affected | Reason/Remarks |
| --- | --- | --- | --- |
| 0.1 | 14/03/2013 | All | First draft version |

# Table of Contents

Chapter 1.	Introduction	7

1.1 Scope of the Document	7

1.2 Documents	7

1.2.1	Applicable Documents	7

1.2.2	Reference Documents	7

1.3 Acronyms and Abbreviations	7

1.3.1	Acronyms	7

Chapter 2.	General Requirements	10

2.1 General characteristics of the Fabry-Pérot Calibrator	10

2.2 Dimensioning of the Fabry-Pérot etalon	11

Chapter 3.	Technical Requirements	13

3.1 Operational conditions	13

3.1.1	Atmospheric pressure	13

3.1.2	Temperature	13

3.1.3	Relative humidity	13

3.2 Performance requirements	13

3.2.1	Optical requirements	13

3.2.2	Fabry-Perot parameters	13

3.2.3	Wavelengths stability	13

3.2.4	Line shape and line shape stability	14

3.2.5	Vacuum	14

3.2.6	Flux	14

3.2.7	Monitoring	14

Chapter 4.	Design architecture	15

4.1 Product Tree of the RV-reference module	15

4.2 Primary source and spectral flux	16

4.3 Etalon parameters	17

4.4 Thermal control	18

4.5 Pressure sensor and controller	18

4.6 Summary of hardware devices	19

Chapter 5.	Mechanical design	20

5.1 Overview	20

5.2 The vacuum tank	23

5.3 Optical feed-through	23

5.4 Fabry-Perot assembly	23

5.5 The entrance fiber-holder spider	24

5.6 The parabolic mirror mounting	24

5.7 The etalon mounting	25

5.8 The exit fiber-holder spider	26

Chapter 6.	Interfaces	27

6.1 Optical interface	27

6.2 Mecanical interface	27

6.3 Control system software interface	27

6.4 Hardware interfaces	27

Chapter 7.	Performance Analysis	28

Chapter 8.	Compliance matrix	29

Chapter 9.	Appendix	30

# List of Figures

Figure 1: Functional diagram of the RV reference unit	15

Figure 2: EQ99 light source and lamp controller	16

Figure 3: Spectral power distribution of the LDLS source.	17

Figure 4: Internal transmission of the FPC fibers	17

Figure 5: Internal transmission of Infrasil Fused Silica	17

Figure 6: Lakeshore 335 Temperature Controller	18

Figure 7: MPT vacuum gauge (right) and its DPG109 controller (left)	18

Figure 8: FPC hardware devices	19

Figure 9: Section view of the Fabry-Perot	22

Figure 10: Ouside view of the vacuum vessel	23

Figure 11: Fiber-optical vacuum feed-though	23

Figure 12: General view of the FPC opto-mechanical assembly	24

Figure 13: View of the upper spider and fiber support	24

Figure 14: Section view of the parabolic mirror and its support	25

Figure 15: Section view of the etalon and its support	25

Figure 16: Section view of the upper spider and fiber support	26

# List of Tables

Table 2: Requirements towards the FPC	13

Table 1: FPC hardware devices	23

Table 2: Compliance matrix	41

# Introduction

## Scope of the Document

To be integrated elsewhere …

## Documents

The applicable and reference documents are listed below:

### Applicable Documents

| AD-1 | ESPRESSO Statement of Work | VLT-SOW-ESO-13520-5059 | 1 | 01.02.2011 |
| --- | --- | --- | --- | --- |
| AD-2 | ESPRESSO Technical Specifications | VLT-SPE-ESO-13520-4633 | 3 | 01.02.2011 |

### Reference Documents

| RD-1 |  |  |  |  |
| --- | --- | --- | --- | --- |

## Acronyms and Abbreviations

### Acronyms

AD	Applicable Document

ADP	Acceptance data package

AIV	Assembly, Integration and Verification

AT	Acceptance Test

BDR	Baseline Design Review

CCB	Configuration control board

CI	Configuration Item

CIDL	Configuration Items Data List

CIL	Configuration Item List

CRB	Change Request Board

CRE	Change Request

DCL	Declared Component List

DIN	Deutsche Industrie Norm

DMPPL	Declared Materials, Parts and Process List

DRB	Delivery Review Board

ECSS	European Cooperation for Space Standardisation

EEE	Electrical electronic electromechanical components

E-ELT	European Extremely Large Telescope

EN	Euro Norm

ESO	European Southern Observatory

ESPRESSO	Echelle Spectrograph for Rocky Exoplanets and Stable Spectroscopic Observations

FAC	Final Acceptance Chile

FDR	Final Design Review

FMEA	Failure Mode and Effect Analyses

FMECA	Failure Modes Effects and Criticality Analysis

FPC	Fabry-Perot Calibrator

FTA	Failure Tree Analyses

HW	Hardware

I/F	Interface

ICD	Interface Control Document

ISO	International Organisation for Standardisation

IST	Integrated System Test

IV	Integration and Verification

MAIT	Manufacturing, Assembly, Integration, Test

MAIV	Manufacturing, Assembly, Integration and Verification

MAV	Manufacturing, Assembly and Verification

MMPP	Materials, Mechanical Parts and Processes

MoM	Minutes of Meeting

MRB	Material Review Board

MTBF	Mean Time Between Failures

MTTF	Mean Time To Failure

MTTR	Mean Time To Repair

N/A	Not Applicable

NCR	Non Conformance Report

NRB	Non Conformance Review Board

PAC	Provisional Acceptance Chile

PAE	Provisional Acceptance Europe

PAV	Procurement, Assembly and Verification

PDR	Preliminary Design Review

PI	Principal Investigator

PM	Project Manager

QA	Quality Assurance

RAMS	Reliability Availability Maintainability Safety

RD	Reference Document

RfW	Request for Waiver

S/S	Subsystem

SOW	Statement Of Work

SW	Software

TBC	To Be Confirmed

TBD	To Be Defined/To Be Developed

UT	Unit Telescope (8.2 meter telescope at Paranal)

VLT	Very Large Telescope

VM	Verification Matrix

WBS	Work Breakdown Structure

WP	Work Package

# General Requirements

## General characteristics of the Fabry-Pérot Calibrator

The FPC is designed as a fall-back solution in case of an LFC of suitable properties would not be available by the end of the ESPRESSO project or beginning of operations. The purpose of the Fabry-Pérot Calibrator FPC is to provide a large number of ultra-stable lines in the wavelength range of the ESPRESSO spectrograph. The FPC shall ensure ‘at least’ adequate short-term stability, i.e. better than 10 cm/s during an observing night. Since the FPC system is fully passive, no guarantee can be given a priori that the stability of 10 cm/s is achieved over long term or even years. It will therefore be operated in addition to an absolute reference, so a ThAr lamp or a stabilized laser to which it will be referenced periodically.

We remind here the properties desired for this module:

- The calibration source must cover the full wavelength domain of the spectrograph.

- The line position and shape repeatability of all lines must be better than 0.1 m/s over the instrument life-time.

- The lines are not resolved by the spectrograph, since otherwise information is lost, respectively we do not take full advantage of the spectral resolution of the instrument. The resolution of the spectrograph is 120’000 in its standard mode, limited by the illuminating fiber size.

- The line distance is perfectly stable and analytically defined.

- The line distance must be minimum 2 and maximum 7 FWHMs of the spectrograph IP.

- The relative intensity of any neighbouring lines must be stable at 10% over time.

- The dynamic range of line intensities over one order must be smaller than a factor 4.

As a first goal, the system shall be made sufficiently stable to be used as a relative reference during an astronomical observing night of typically 12 hours. The module will possibly need to be re-calibrated every night, e.g. with a ThAr lamp or a LFC, for absolute referencing, but will serve as perfect simultaneous reference within the night. The second objective will be to extend the time during which the module remains stable. Ideally, the module will be intrinsically so stable, that re-calibration will not be needed over large time scales (years).

Reflection is obtained by a multi-layer thin film coating of the inner optical surfaces. The following aspect must be considered: The apparent separation of the two spacers depends on the effective penetration depths of the wave on the reflective surface.

- Because of the varying real and imaginary part of the refractive index as a function of wavelength, the effective FP separation will vary with wavelength. This wavelength dependence must be avoided as far as possible, in particular, a temperature or aging effect must be avoided. If this wavelength dependence cannot be avoided completely, it should be at least very “smooth”.

- Thermal effects in the coating will affect the way the etalon behaves: The dielectric stack will have a certain thickness. When the ambient temperature changes, the stack will expand (CTE of some 10-6) and this will change the effective gap of the etalon. With a gap of 7 mm, for instance, and a dielectric stack of 4 µm thick with a CTE of 7 10-6, the variation of the apparent separation will be about 8 10-9 / K. In the same way, the index of refraction will change with temperature and affect the apparent penetration depth of the optical wave in the coating.

- Aging of Zerodur has been studied in metrology labs. Schott data indicates shrinkage of 10-7 per year. However, the experience on the HARPS FP shows that the effect of aging is well below the estimated value. In alternative to Zerodur, ULE shall be used.

When making trade-off, the following order of priorities shall be considered:

- Relative wavelength stability over temperature and time (between two neighboring lines)

- Wavelength stability over temperature and time

- Equidistance of transmitted wavelengths

- Uniformity of line intensity and line shape and width

## Dimensioning of the Fabry-Pérot etalon

The primary parameters for which the Fabry-Pérot etalon has to be dimensioned are :

Variables:

- Spectrograph resolution RS

- Shortest wavelength λB

- Longest wavelength λR

Requirements for an optimum use of the lines are:

- The line separation is minimum on the blue side, because of the FP equation. The separation must be chosen such that the lines can be resolved by the spectrograph:
The peak separation at all wavelength must be > 3 FWHM: Δλ≥3∙λRS 	(R1)

- The line separation is highest on the red side. At a given Finesse, the FP lines become wider there. It shall be avoided, that the spectrograph resolves the lines:
The FP line must be under-resolved by a factor of 2/3 : δλ≤2/3∙λRS	(R2)

Let us now first derive the required FP. Given the gap D of the etalon, the resonance of the FP will be achieved whenever	2D=mλ 	(1).

Therefore, the separation between the 2 bluest peaks will be, using requirement (R2):

Δλ=2Dm-1-2Dm=2Dmm-1≥3∙λRS 	(2)

Let's replace m in the last inequality by using equation (1),

2Dmm-1≅2Dm2=λ2/2D≥3∙λRS	 (3)

and solve for D. In order to be true for all wavelengths, the relation must me satisfied for the shortest wavelength λ=λB, such that we get   D=λB∙RS/6 	(P1)

Now let's derive the required finesse F of the etalon: The Finesse is defined as the ratio of the line separation to the FWHM:	F=Δλ/δλ	 (4)

We can express again Δλ as a function of the gap and the FP order

Δλ=2D/m(m-1)≅2D/m2 	(5)

and introduce equation (5) and requirement (R2) into equation (4)

F=Δλδλ≥2D3RSm22λ=3DRSm2λ=3DRSλ24D2λ=3RSλ4D 	(6)

where in the step before the last we used the etalon equation (1). Since this condition must apply at all wavelengths, we have to choose F for the longest wavelength and substitute D using (P1):

F=3RSλR4D=18∙λR4λB=9λR2λB	(P2)

Let's now look at the etalon parameters from the instrument figures: RS = 120'000, λB=380nm, λR=780nm we find

D = 7.6 mm

F = 9.2

# Technical Requirements

## Operational conditions

The FPC will be operated in the CCL at the Paranal Observatory, Chile. It will be located within 2T. The FPC will not be exposed to any motion, gravity, vibrations, etc.

### Atmospheric pressure

Atmospheric pressure is 760 mbar in average. Given the high ambient pressure variations the FPC etalon will be operated in vacuum. A pressure value of 0 mBar must be considered for the optical computations.

### Temperature

The temperature of the environment is typically of TBD ± TBD°C.

### Relative humidity

The system is working normally in a controlled environment with relative humidity always lower than 100%. During manufacturing, transportation, installation, and maintenance the humidity can attain peeks close to 100%.

## Performance requirements

### Etalon requirements

Operational wavelength range:	380 nm - 780 nm

Total transmittance at peak:	T > 10% 	in the wavelength range

Transmittance uniformity	Tmax/Tmin < 2	in the wavelength range

Transmittance variations	dT/d< 2% per nm	in order not to distort the

Lines by more than 10 cm/s

### Fabry-Perot parameters

Effective total Finesse:	10 < F < 12		in the wavelength range

Fabry-Perot spacing:	D =  7.6 mm  0.0005 mm

In order to achieve a total (reflectivity-limited) Finesse of F = 10, the various Finesse contributions must be:

Reflectivity Finesse: 	FR≥54∙10=12.5

Aperture Finesse:	Fθ≥2.35∙F=23.5

Defect Finesse: 	FD≥2∙2.35∙F=33.2

Parallelism Finesse:	FP≥2∙2.35∙F=33.2

### Delivered flux and flux homogeneity

The FPC shall deliver sufficient flux at peak wavelength within the étendue accepted by the spectropgraph, such that a global photon-noise precision of better than 5 cm/s is achieved in less than 20 s.

The white-light sources should be selected and/or balanced in a way that the ratio between maximum and minimum counts per extracted pixels measured all over the spectrum must not exceed a factor of 5 (dynamic range).

### Wavelengths stability

The Fabry-Perot transmitted wavelengths must remain stable with time. The Fabry-Perot shall be placed in vacuum, and the required temperature stability specified. Changes of the transmitted wavelengths must be minimized.

Wavelength repeatability:	d/dt < 2 10-10 7 m/s	at any wavelength

required: during at least 12 hours

goal: during 10 years

Absolute wavelength:	λ - λ0 	3 x -8λ (10 m/s)	at any wavelength

The absolute wavelength of a line is defined on the centroid of the transmitted FP line with respect to the expected wavelength given by the nominal etalon spacing. Wavelength repeatability indicates the stability of the centroid of a line.

The Fabry-Perot must be optimized for minimum sensitivity of the transmitted wavelengths with regard to temperature variations. If the final design is intrinsically stable it relaxes the contraints on the thermal control stability.

### Line shape and line shape stability

The shape of the transmitted FP lines (intensity, symmetry, width) shall be as uniform as possible. The line shape shall be as stable as possible. It is therefore suggested that the etalon finesse is determined by the reflectivity finesse, which is the most insensitive to geometrical aspects. However, in order to make the transmitted lines 'equidistant' to a level described by the 'absolute wavelength' requirement, the phase of the etalon transmittance function, and thus the reflectance of the individual mirrors, shall be kept constant as a function of wavelength, if possible. Since this is very difficult, the plan will be to calibrate the phase variations with an FTS at the beginning of the lifetime of the FPC.

### Vacuum

From the stability requirements we deduce that the air pressure around the FP etalon must be kept stable within 10-3 mbar during 12 hours. The requirement shall be therefore formulated as:

dp/dt < 2 x 10-3 mbar per day.

### Monitoring

The temperature of the FPC, and in particular the etalon, must be monitored and controlled continuously with a resolution of 0.001 K, a short-term repeatability of 0.002 K over time scales of 1 months and of 0.025 K over five years, and an absolute accuracy of 0.1 K. The monitoring must have a data link to the ESPRESSO control system.

The air pressure around the etalon must be monitored continuously with a resolution and precision of 0.0001 mbar, and an absolute accuracy of 0.001 mbar. The monitoring must have a data link to the ESPRESSO control software.

### Summary of performances requirements towards the FPC

In the following Table 1 we summarize the relevant technical requirements towards the FPC.

Table 1: Requirements towards the FPC

| Item | Requirement | Comment |
| --- | --- | --- |
| Wavelength coverage | 380 – 780 nm | Wavelength range of ESPRESSO |
| System transmittance | > 10% | In the wavelength range |
| Transmittance uniformity | Tmax/Tmin < 2 | In the wavelength range |
| Transmittance variations | dT/d< 2% per nm |  |
| Etalon total Finesse | 10 < F < 12 | In the wavelength range |
| Etalon spacing | D = 7.6 mm  0.0005 mm |  |
| Photon noise | < 5 cm/s | Global precision achieved in a single exposure |
| Flux homogeneity | Fmax/Fmin < 5 | In the wavelength range |
| Line width | < 2/3 of spectral element | A spectral element is the wavelength divided by the spectral resolution |
| Line separation | 2-7 x spectral element |  |
| Short-term repeatability | d/dt < 2 10-10 7 m/s | Over 12 hours |
| Long-term repeatability | d/dt < 2 10-10 7 m/s | Over 20 years (goal) |
| Local wavelength accuracy | λ - λ0 3 x -8λ (10 m/s) | For any FPC line |
| Line shape stability | Finesse determined by FR |  |
| Local wavelength accuracy | λ - λ0 3 x -8λ (10 m/s) | For any FPC line |

# Design architecture

The proposed FPC is derived from a new concept that has been successfully evaluated on HARPS and HARPS-N: The stability reference for the radial velocity is essentially a Fabry-Perot etalon housed in a temperature controlled vacuum enclosure. The etalon is fibre-fed by a bright white lamp. A symmetrical set-up of parabolas couples the input to the exit fibers, with the etalon located in the collimated beam between the two parabolas, making the design achromatic.

Figure 1: Functional diagram of the RV reference unit

## Product Tree of the RV-reference module

- White source

- LDLS source with SMA or FC fiber connection

- Balancing filter (color correction)

- 1-m fiber 230 µm SMA/FC-FC with metallic protection jacket

- Optics

- 1 fiber vacuum feedthrough

- 1 thin input fiber FC-FC of 230 µm dia. without jacket

- 2 protected-silver coated parabolic mirrors with f = 100 mm

- 1 etalon

- 1 thick output fiber FC-FC of 600 µm without jacket

- FP-etalon assembly (structure according to drawing)

- Vacuum tank

- Electroplated aluminum body

- Dismountable cover

- DN-40 pumping flange

- 2 fiber-feedthrough mounting flanges

- 1 electrical feedthrough mounting flange

- Temperature control device

- 1 Lakeshore 335 controller

- 1 connect cable for heater and sensors

- 2 wired Si-diodes temperature sensors

- 2 wired heating foils

- Vacuum monitoring device.

- Digiline pressure sensor

- Connection cable

- DPG-109 pressure-gauge controller

## Primary source and spectral flux

For irradiance reasons, spectral flatness and MTBF, we would like to use a innovative source called “laser driven light source” EQ-99FC Plus from the company Energetiq. This source us presently being used in integration on HARPS North with convincing results. The light source possesses a complete remote control system for control via software. However, the lamp will be used in stand-alone mode, i.e. it will manually switched ON and kept ON continuously for FPC stability reasons. The lifetime of such a lamp is longer than 20’000 hours. A LRU spare shall be procured for immediate exchange in case of failure. For the fibers standard Ceramoptec octagonal fibers with high internal transmittance will be chosen.

Figure 2: EQ99 light source and lamp controller

Figure 3: Spectral power distribution of the LDLS source.

Figure 4: Internal transmission of the FPC fibers

In order to reduce the heat-up of the FP etalon by the source, we will introduce in the laser source a heat-blocking filter TKG-5253 by UQG Optics with the transmittance curve shown is Figure 5. As a function of the obtained chromatic distribution, a balancing filter will be added. For the moment no choice has been made, since a) a balancing filter may not be necessary, given the high blue flux of the Fabry-Pérot, and b) we prefer in any case to see the resulting chromatic flux distribution. A filter can be inserted easily at anytime an a appropriate location will be reserved for it.

Figure 5: Heat-blocking filter TKG – 5253 transmittance

## Etalon parameters

Given the gap and Finesse requirement, the etalon will be designed as follows:

Gap: D = 7.6 mm

Clear diameter D = 40 mm (design), focal length of collimator f = 100 mm

Reflectivity Finesse: 	FR = 12.5	R() = const = 78%

Aperture Finesse:	F = 23	Fiber diameter d < 300 microns @ fcoll = 100 mm

Defect Finesse: 	FD = 33	Mirror wavefront better than l/50 P-V

Parallelism Finesse:	FP = 33	Departure from parallelism < 10 nm -> 0.25 10-7 rad

The etalon plates will be manufactured form fused silica blanks, most likely INFRASIL. The internal fibers being very short, we propose to use Ceramoptec WF fibers as well.

Figure 6: Internal transmission of Infrasil Fused Silica

## Thermal control

From our experience with previous similar systems, the thermal control of the etalon structure to a couple of mK rms should be a straightforward affair. It will be based on a Lakeshore 335 temperature controller using silicon diodes as sensors and one or two thermal pads as heating elements.

Figure 7: Lakeshore 335 Temperature Controller

## Pressure sensor and controller

The pressure of the FPC and thus around the etalon will be measured by means of an MPT100 Pfeiffer gauge attached to the vacuum tank. A DPG109 controller, the one used for the Vacuum System, will read out the sensor. The DPG controller will be read out through the ESPRESSO ICS.

Figure 8: MPT vacuum gauge (right) and its DPG109 controller (left)

## Summary of hardware devices

Table 1: FPC hardware devices

| Type | Manufacture | Reference | Interface | Notes |
| --- | --- | --- | --- | --- |
| Temperature controller | LAKESHORE | 335 | USB or Ethernet |  |
| Temperature sensor | LAKESHORE | Si diodes DT-670 |  | Silicon diode |
| Pressure controller | PFIEFFER | DPG 109 | RS-232 (DB9) | 9 inputs |
| Pressure gauge | PFIEFFER | MPT100 |  |  |
| Lamp controller EQ99 Plus | Energetix |  |  |  |

# Mechanical design

## Overview

The vacuum tank is a simple cylinder of electro-polished aluminum with 2 aluminum covers. It contains the actual FPC and its etalon, which is sketched in the following. While the first model of the FPC had a structure made of optically contacted Zerodur cylinders, a sensitivity analysis demonstrated later that this was not necessary. The mechanical structure of the second model (Harps North) is more classical and a less expensive assembly of aluminum elements holding the mirrors, the etalon and the fiber connectors.

Parabolic mirror

Adjustable spider in X,Y,Z

Output fibre

Etalon

Input fibre

Parabolic mirror

Adjustable spider in Z

Aluminium body

Vacuum vessel

Input feed through

Output feed through

Figure 9: Section view of the Fabry-Perot

Description of the light path

The light coming from the front-end enters the system by a vacuum feed-through (1), then goes through an optical fibre (2), fixed at one end to a spider (3). The light coming out of the fibre is projected against a parabolic mirror (4) which sends back the collimated beam trough the etalon (5) and against the second parabolic mirror (6). This mirror focuses again the beam and the light is finally collected by the output fibre and goes out trough the second vacuum feed-through.

4

6

5

3

2

1

Light output

Light input

## The vacuum tank

The vacuum vessel is a hermetic chamber in which the pressure is maintained at 1E-3 mbar and the temperature is controlled. It is made out of Aluminum and is composed by 3 parts:

- A centre body (1): it is a tube with 2 welded flanges at each end.

- A blank flange (2) on the bottom side.

- A flange with feed-through (3) on the upper side. 2 optical feed-through (4) and 1 electrical feed-through (5). There is also a DN 40 ISO-K half nipple flange welded on for the pumping connection.

2 seals insure the airtight and 8 clamps maintain all parts together.

Figure 10: Ouside view of the vacuum vessel

4

2

5

3

1

## Optical feed-through

It is an off-the-shelf part composed by a flange and 2 FC receptacles at both ends. An optical fibre passes through a soft sealant material, which is compressed by the compressive end-cap and internal squeezer of the feed-through.

Figure 11: Fiber-optical vacuum feed-though

## Fabry-Perot assembly

It is a stand-alone system placed in the vacuum vessel and fixed to the top flange in 3 points. Some insulated housings avoids thermal transmission from the outside.

Figure 12: General view of the FPC opto-mechanical assembly

Vacuum vessel top flange

First spider

Input optical fibre

Second spider

Parabolic mirror

Etalon

Output optical fibre

Parabolic mirror

Insulated housings

## The entrance fiber-holder spider

This is the support for the input optical fibre connector (1). This spider (2) is adjustable only in vertical (Z) direction by shimming spacers (2).

Figure 13: View of the upper spider and fiber support

3

2

1

## The parabolic mirror mounting

These are the two mirrors (1) at each end of the system. The first one collects the beam coming from the input fibre and collimates it back through the etalon. The second one collect the light coming from the etalon and focus it back to the output fibre. These mirrors have a reference surface (2) on the same side as the parabola (3). A Viton® O-ring (4) compressed by an aluminium ring (5) maintain the mirror in position. 3 spacers (6) avoid a contact between the mirror (1) and the aluminium ring (5).

Figure 14: Section view of the parabolic mirror and its support

6

3

2

5

4

1

## The etalon mounting

The etalon (1) is the main optical part. At its base, it has a flange to fix it. 3 brackets (2) and a Viton® O-ring (3), used as a sprig, maintain the etalon in position on the supporting flange (4).

Figure 15: Section view of the etalon and its support

4

2

3

1

## The exit fiber-holder spider

This is the support for the output optical fibre connector (1). This spider (2) is used for the final optical adjustment. It is adjustable in all directions (X,Y,Z): 4 pushing screws (3) for X,Y direction and 3 spacers (4) to shim for the vertical direction. The 4 screws are removed after adjustment. To insure an accurate displacement in X and Y direction, an intermediate ring (5) is placed between the supporting flange and the spider. In this intermediate ring are machined grooves (6) in the 2 directions.

Figure 16: Section view of the upper spider and fiber support

6

4

3

2

1

5

# Interfaces

## Optical interface

The FPC will have only one optical interface to the calibration unit. It will provide a male FC connector holding a 600 µm fiber to be connected at the input port of the calibration unit.

## Mechanical interface

The opto-mechanical part of the FPC will be housed in a 0.6m x 0.6m x 1m enclosure of 60 x 60 x 80 cm. This enclosure is mainly filled with insulating foam and hosts the vacuum tank in the middle. This enclosure weights about 25kg and must be located in a stable thermal environnment (2T).

The primary light source will be packaged in a 6U 19” rack. The lamp control system will be packaged in a 3U 19” crate, as well as the Lakeshore temperature controller and the vacuum gauge controller

The Fabry-Pérot will be attached to the pumping system of the vacuum system for initial pumping and possible periodic pumping. Normally, no pumping will be performed on a regular basis, however, given the fact that the pressure increase is sufficiently small to avoid nightly drifts larger than the maximum requird10 cm/s, and, becse long-term drifts will be calibrated with an external absolute calibration reference

## Control system software interface

No active control is to be provided. The FPC will only be connected ‘passively’ to the ICS for monitoring of:

- Lakeshore 335 temperature controller for the two temperatures of the FPC

- Pressure DPG109 for the readout of the pressure gauge. The pressure gauge will be actually be connected to the Vacuum System pressure controller.

## Hardware interfaces

The FPC is a stand-alone and passive unit. No remote power control is required since it will be continuously operated. The unit is turned ON and turned OFF manually, whenever necessary. The controllers are installed in standard 19” rack within a small cabinet. A single standard 230V50Hz UPS power line must be delivered to the cabinet

The Lakeshore 335 is accessible via an Ethernet or serial port

The Pfeiffer DPG109 pressure controller of the Vacuum System will be used to read-out also the FPC pressure gauge. The DPG is controller through a serial port.

# Performance Analysis

In the following we shall present the results obtained with a FPC already produced for HARPS and HARPS-N (Figure 18). These results replace any analysis, since it provides real-life and assessed performances. Figure 18 shows the Etalon manufactured by ICOS and used by HARPS and HARPS-N.

Figure 17: FPC currently in use on HARPS-N during the tests (left) and in its isolation box (right)

Figure 18: Fabry-Pérot etalons produced by ICOS. Left: Pure Zerodur-spacers etalon of low CTE. Right: Compensated spacers for theoretical zero CTE

## Spectral range and energy distribution

The spectrum of the FPC system as recorded by the HARPS spectrograph is shown in Figure 19. Only a part of the raw frame is shown, as the whole image would be too large and the various lines would not be resolved individually. One can see the spectra of the two fibers both illuminated by the FPC system spectrum. The wavelength increases from left to right and from bottom to top, direction over which the spectrum is split over several echelle orders. The large gap in the center of the image is due to a physical gap in the CCD mosaic composed of two CCDs. For comparison a frame is shown in which one of the fiber was illuminated with the thorium lamp. Note the richness of the FP spectrum compared to that of the thorium!

Figure 19: Top: Raw HARPS frame of the FPC spectrum illuminating both fibers. Bottom: For comparison, a frame is shown in which one of the fiber was illuminated with the thorium lamp. Note the richness of the FPC spectrum compared to that of the thorium.

The average transmission of the FPC system has been measured by using a low-resolution laboratory spectrograph. The spectrograph recorded the white lamp spectrum alone and the same spectrum filtered by the FP system. The ratio of both spectra, which is basically the average transmission of the FP system including the fibers and the connectors, is shown in Figure 20 on the left side. Down to 450 nm the average transmittance is 10%. In order to determine the transmission at peak, however, this value has to be multiplied by the Finesse of the etalon. As we will see later, the finesse is 4.3; therefore the peak transmission must be of the order of 43% above 450 nm. In we consider that a significant fraction of the transmission losses are due to the two fiber connectors, we conclude that the peak-transmittance of the FP system is well above 50%. The much lower transmittance at short wavelength is due, on the one hand, by the fibers and the optical materials, which have generally lower efficiency towards the blue and, on the other hand, is caused by the FPC-etalon coating, which was not optimized for this wavelength region.

Figure 20: Left: Transmission spectrum of the Fabry-Pérot etalon averaged of a broad band (>> than a typical FSR of the etalon). Right: flux at the peak of a transmission line of the FP when illuminated by the Xe-lamp compared to the tungsten lamp. Measurements have been carried out with HARPS.

The right side of Figure 20 shows the flux level at peak transmission across the FPC spectrum as recorded by HARPS. The indicated flux includes the lamp, the FP system transmittance and the spectrograph efficiency. For comparison, the same curve is shown when using a tungsten lamp. It can be seen that, for the tungsten lamp, the flux level in the 15 bluest orders is too low for them to be used for the wavelength calibration and drift computation, while only a few orders are lost in the case of the Xe-lamp. It must be noted that the curves of the Xe and the tungsten lamp have been normalize, such that only the relative spectral energy distribution is shown. The total flux ratio between the Xe and the tungsten lamp is of the order of 16!

In order to increase lifetime, the Xe-lamp on HARPS has been replaced by a super-continuum laser source at the cost of some flux. The achieved photon-noise precision on HARPS is 5 cm s-1.   On HARPS-N the design was improved and the lamp replaced with an LDLS. An RV precision of about 2 cm s-1 per frame is obtained on a single frame, which is about three time as good as using the ThAr for the simultaneous drift measurement. The flux per extracted pixel at the peak of the echelle order with the FPC system on HARPS-N is shown in Figure 21. Order 1 corresponds to 690 nm, while order 69 is about 385 nm. Thanks to a coarse but ‘wavy’ interference filter, the red-to-blue flux has been well balanced. A part from the four bluest order, the flux dynamics remains well within the specified factor of 5. The fast an unavoidable drop in the extreme blue is due, as mentioned above, to the blue-absorbing coating of the etalon. A more balanced coating shall be designed for the ESPRESSO FPC.

Figure 21: Peak flux of the FPC as a function of echelle order (from 690 nm, left, to 380 nm, right)

## Finesse and spectral resolution

The finesse F =  (distance of two neighboring lines divided by the line width of a line) of the FP etalon has been first measured in laboratory by scanning the etalon’s effective gap. The scanning was achieved by varying the pressure inside the vacuum vessel and consequently between the two mirrors of the etalon. Thus, an effective optical path difference (OPD) is introduced and the transmitted wavelength changed. A commercial HeNe laser of  = 633 nm was used for the line scanning measurement.

Figure 21 shows the transmitted laser flux as a function of effective optical gap variation as computed assuming a linear relationship between pressure and refractive index of air. The red curve corresponds to the Airy function which theoretically describes the transmittance function of the FP etalon with a single effective finesse FE. The measured FE is in our case simply the ratio between the FWHM of a single peak divided by the distance between two neighboring peaks. The so obtained value is FE = 4.1. This value is lower than what expected from the design, which predicted an effective finesse of about 6. However, we have to take into account that the natural line width of a HeNe laser that is expected to be of the order of 3.1 10-6 of its frequency, which would correspond to a resolution R =  = 316’000. At 633 nm, the FP etalon is working in the mode m = 23’064. Thus, the corresponding finesse is FL = RL/m = 13.7. The effective finesse of the FP etalon can then be estimated to , where Fm = 4.1 is the measured finesse. We finally obtain FE = 4.3. It has to be noted that this value is still lower than expected.

Figure 22: Direct measurement of the FP transmittance when varying the optical gap

The left side of Figure 22 shows a zoom on the raw frame of the FP spectrum as recorded by HARPS. The left side shows the corresponding ‘extracted’ spectrum, which represents the flux as a function of CCD pixel (or wavelength). The uniformity of the intensity and the equidistance of the spectral line must be remarked. Figure 23 shows the computed ‘finesse’ as a function of wavelength (or spectral order). The computed finesse does not correspond to the FP finesse, since the lines are convolved with the instrumental profile (IP) which determines the spectral resolution of the HARPS spectrograph to about R = 115’000. The finesse so recorded is rather the ratio of the distance between two neighboring lines and line’s FWHM. If the line width was dominated by the spectrograph’s resolution, we would expect a finesse of 3.1 at 400 nm (order nr. 8) and of 5.3 at 680 nm (order nr. 70). The lower measured value confirms, however, that the effective finesse of the FP etalon must be 4.3 instead of 6, in order to obtain the actual measured finesse.

Figure 23: Zoom on the FP raw spectrum as recorded by the HARPS spectrograph and portion of the extracted 1-D spectrum

Figure 24: Ratio of line separation to line width over the full wavelength range of the HARPS spectrograph. Order nr. 1 corresponds to 383 nm while order nr. 70 is located at about 680 nm. At short wavelengths some noise is introduced by the low flux. Within a single order the dispersion of the measured finesse is due to the varying dispersion of the echelle spectrograph along the order.

## Line positions

For the theoretical Fabry-Pérot the wavelength (or frequency) of a given transmitted peak (order) m is fully determined by a single parameter: the gap D. The corresponding formula is m = 2D/m. Provided that we know the etalon’s gap, we can, from this formula, directly determine the wavelength of a given line by knowing it order number. In practice, however, one does not know the absolute order number of the line, but only a relative numbering by assuming that the peaks must have continuously increasing and discrete order values. We have therefore reversed the problem by using the spectrographs calibration delivered by the thorium source and used it to assign a wavelength to each transmission peak. Using this wavelength and assuming a gap of 2D = 14.6 mm we have been able to determine directly the order number of each transmittance peak appearing in the FP spectrum.

We did not expect the gap D to be exactly the theoretical values, nor to be constant as a function of wavelength due to the varying effective optical depth of the dielectric mirror coatings. Indeed, the computed order numbers were not all perfectly discrete. Nevertheless, the difference to the discrete continuous number is always below 0.25. From this we can conclude that by rounding the obtained order number to the closest integer value we unambiguously identify the real order number. Using this fact, we have then used the so computed discrete order number and the wavelength assigned from the thorium calibration to determine the effective gap D = D() of the etalon. The result is shown in Figure 24. Note the mechanical gap of the etalon is of 7.3 mm and that in the figure the optical path difference 2D is actually shown.

Figure 25: Effective etalon gap 2D() as a function of wavelength and is difference from the nominal value of 14.6 mm

It appears that the ‘manufactured’ mechanical gap of D = 7.3 mm is fully compliant with the measured optical path difference of 2D = 14.6 mm within a quarter of wavelength, which confirms that the order numbering can be assigned unambiguously by assuming the theoretical gap value. If a wrong value D had been assumed, the D() would show a slope of at least one wavelength over the plotted spectral range.

Nevertheless, D() seems to vary significantly. The measured changes correspond to about 10-5 or, expressed in radial velocities, of the order of 3 km s-1. We assume here that D() is constant in time and that thus the presented fact does not have any impact on ability of the FP system to track instrument drifts. However, we also conclude in order to use the FPC system for absolute wavelength calibrations to better that 10-5, D() of the Fabry-Pérot must be calibrated by another mean (e.g. laser frequency comb or thorium lamp).

Although D() seems well determined we have to remark that short-scale patterns with amplitude of the order of 50 m s-1 have been recorded in D(). At first glance the patters seems to be linked to the FP system itself (the pattern is the same when comparing wavelengths present in two different orders), and should be calibrated..

We conclude here that the FP etalon behaves as expected in terms of spectral performances. The etalon gap has been manufactured (and measured) with an amazing accuracy. We confirm that the effective gap D() varies with wavelength. However, we will see later that this gap D() remains constant with time and can be thus calibrated down to photon noise precision.

## Photonic Doppler precision content

The ability of determining the position of a spectrum on the CCD (line position expressed in pixel) and associate it with a given wavelength depends on the width and the intensity of the line. In general one can say that the fundamentally attainable precision is proportional to the square root of the number of lines in the spectral domain and their intensity, and inversely proportional to their width. In HARPS, the formulae described by Bouchy & Pepe (2001) are employed to measure the fundamental (photon-noise limited) precision content of a spectrum. Figure 21 shows the attainable precision using the FPC system as a function of the square root of the relative flux measured on the detector. The flux was varied either by changing the variable ND-filter in front of the calibration lamp or by varying the exposure time. For comparison, also the results obtained with the hollow cathode thorium lamp are shown. In both cases it could be demonstrated that the relationship is perfectly linear as expected from pure photon noise, and that a precision of about 2 cm s-1 can be attained on a single spectrum.

Figure 26: Fundamental (photon-noise limited) precision obtained on a single FP spectrum. The results from the thorium lamp are shown for comparison.

As mentioned above, the flux level can be varied either by increasing the light intensity or the exposure time. In the present HARPS-N system the thorium lamp is tuned in a way such to balance lifetime and flux. In an exposure of 20 s, which is the (minimum) nominal exposure time for a wavelength-calibration exposure, the attained precision level is of the order of 2 cm s-1.

## Stability measurements

### Drift computation and use of new algorithm

In order to use the FPC in a standard way on HARPS, the DRS (pipeline) had to be adapted and some functionality added. In particular, one had to add the possibility of treating frames using the FPC on fiber B for both calibration and scientific (target) exposures. The recognition and the correct handling of the various cases had to be implemented. Also, for the measurement of the drift, several new algorithms have been tested and implemented.

It is important to note that these developments have been also used for the simultaneous development of the LFC by ESO and MPQ, since the LFC as well had to be tested on HARPS. The spectra of the LFC and the FPC are actually qualitatively identical, and the same data reduction can be used in both cases. These concepts have been fully integrated in the ESPRESSO pipeline and apply in an identical way for the Laser-Frequency Comb and the Fabry-Pérot.

### Short-term stability

We have tested the short-term stability by performing long series of ThAr-FP exposures (ThAr on fiber A, FPC on fiber B) and by computing the respective drifts independently. Figure 19 shows the drift from one exposure to the other as computed by the ThAr lamp and the FPC, as well as the difference of both to remove possible instrumental drifts.

For the ThAr alone we obtain a dispersion of about 0.19 m s-1 rms which should be compared to about 0.09 m s-1 pure photon-noise error. For the FPC we obtain 0.176 m s-1 rms dispersion, to be compared to 0.06 m s-1 photon noise. For the difference of both, we end up with a photon noise of 0.11 m s-1 and a measured dispersion of 0.14 m s-1. The dispersion of the dirrence is well below the square root of the quadratic sum of the individual drift measurements of 0.26 m s-1, demonstrating that both sources track a ‘real’ instrumental’ drift. On the other hand, the difference between the photon noise and the measured dispersion leaves room for a 0.08 m s-1 additional dispersion, which is not ‘seen’ by at least one of the fibers (or better sources).

Figure 27: Relative drift measurements over 7 hours to test the short-term performances of the FPC

The plot in Figure 20 shows the same data, but this time in absolute terms, i.e. it indicates the absolute drift for FP and TH with respect to the very first exposure of the series. It should be remarked that in general the FP and the TH follow the same ‘high-frequency’ behavior, tracking real instrumental drifts. This is demonstrated by the fact that the dispersion of the difference is well below the dispersion of the individual measurements. On the other hand, the plot shows also that one can have differential drifts of the order of 0.9 m s-1 P-V or 0.23 m s-1 rms if no recalibration is performed during every night. Again, this is a worst case limit if we assume that the differential drift between TH and FP is fully due to the FP. As it will be demonstrated in the following, the short-term stability oft he FPC is well within specs.

Figure 28: Absolute drift measurements over 7 hours to test the short-term performances of the FPC

### Long-term stability

The long-term stability was tested in a similar way as the short-term stability using HARPS@ESO data. Frames were taken in routine operation almost every day. The idea was to monitor possible long-term drifts of the FPC with respect to the ThAr calibration lamp.

Figure 21 shows the results. Over the 6 months of operations in 2012, the FPC slowly drifted by a total of 11.6 m s-1, assuming that the ThAr lamp has no internal drift. The measured drift is mostly due to the fact that the pressure in the FPC increases slightly with time, since it is not continuously pumped. This sequence shows nevertheless in a clear and impressive way that the long-term drift of the FPC is of about 0.063 m s-1 per day in average. During an observing night, between the calibration and the scientific exposures, we should not expect a drift larger than 0.05 cm s-1, if the daily calibrations are done at the beginning of the night.

Figure 29: Drift measurements over 180 days to test the long-term performances of the FPC

### Wavelength dispersion and local accuracy

The effective etalon gap varies as a function of wavelength due to the fact that the penetration depth in the coating depends on the wavelength. By using the ThAr calibration of the FP frames. We have repeated this measurement in order to verify whether the etalon characteristic, in particular the coatings, remain constant with time. Figure 30 shows the effective gap (twice the physical gap) as a function of wavelength. The two curves and the respective zoom show a measurement taken with the continuum laser source of HARPS in 2011 and another in 2012 more than one year later. Both plots show that the two curve follow each other to a large degree. The absolute variation of 0.3 µm corresponds to a ‘velocity’ of  6 km s-1. The difference between the two curve is much smaller and basically an offset. It is mainly due to a pressure change in the FPC system which influences the effective gap of the etalon. Small ‘coherent’ variations can be seen on short scale. They correspond to about 0.003 nm or 60 m s-1.

Figure 30: Dispersion law of the FPC (expressed in effective etalon gap) recorded over two seasons separated by more than one year. Left: full wavelength range of HARPS.  Right: Zoom over a single order.

Figure 31 shows the difference in dispersion between the two seasons separated by more than one year. The variation is expressed in terms of effective etalon gap variation. As it can be seen, the variation is constant with wavelength. The general offset of 0.65 nm (= 13.4 m s-1) can be explained by the differences in pressure inside the etalon dewar between the two measurements. It interesting to note that the two curves show the same low-frequency behavior, confirming that the coatings characteristics remain unchanged to the level given by the measurement noise of each single line.  It should be noted that the difference of the dispersions removes the ‘coherent’ wiggles seen in the absolute dispersion curve shown in the previous figure. The remaining difference of typical ±1 nm or about ±20 m s-1 is probably simply due to ‘photon noise’ when determining the line position on one single frame. Indeed, a back of the envelope estimation of the centroid error obtained on a line of 5000 electrons/extracted pixel, a typical value for the HARPS FPC frames, leads to an error of 20 m s-1. On the HARPS-N FPC the flux is about a factor of 6 higher, such that we can expect that the FPC lines are determined with a precision of about 2.5 times better, i.e. to a typical precision of 8 m s-1.

Figure 31: Dispersion variation (expressed in difference of effective etalon gap) between two seasons separated by one year as a function of wavelength. Left: full wavelength range of HARPS.  Right: Zoom over a single order.

## On-sky tests and operations

### Description of operations

In order to operate the FPC in a standard way, we have implemented a operational scheme in an identical way as when the ThAr lamp is used. Templates had to be added and the DRS modified in order to do so. In the following we describe the operations step-by-step:

- During the calibrations, a wavelength-calibration exposure is taken where fiber A, the object fiber, is illuminated with the lamp ThAr 1 and fiber B, the reference fiber, with the FPC. This exposure is produced by calling the calibration Template ‘HARPS_ech_cal_WAVE’ with the fiber B parameter set to ‘FP’ by default.

- The DRS recognizes this frame as such and performs extraction and wavelength calibration on fiber A and extraction and referencing on fiber B. The result is stored in the calibration database calib_DB under the reference ‘wave’. It shall be noted that for the wavelength calibration of fiber B the latest ‘THA2’ exposure (both fibers illuminated with ThAr) is used.

- During the scientific observation fiber A, the object fiber, is illuminated with the astronomical target and fiber B, the reference fiber, with the FPC. This exposure is produced by calling the acquisition template ‘HARPS_ech_acq_WAVEsimult’ with the fiber B parameter set to ‘FP’ by default.

- The DRS recognizes this frame as such and performs extraction and stellar velocity computation on fiber A and extraction and drift computation on fiber B. For the latter, the reduction recipe looks into the calib_DB for the latest ‘WAVE’ calibration with ‘FP’ on fiber B and computes the differential drift between the present exposure and the reference exposure.

This scheme has been implemented and is fully operation. Two ‘HARPS_ech_cal_WAVE’ using ‘FP’ on fiber B have been included in the Standard-RV calibration set and are executed every night on HARPS. As mentioned above a similar scheme is foreseen for ESPRESSO.

### Observations of known standards

In order to test the on sky performance we have observed, during the ‘Sousa’ mission in March 2011, two well-known targets. Both targets are part of the ‘HARPS-Upgrade GTO’ program by Pepe et al. In order to compare the radial-velocity results, we have observed both targets in their standard modes (usually used modes). The first target, HD 85512, is a moderately bright stable star with one known planet. Usually, this target is observed in ‘thosimult’ mode. The second star, HD 128621 or alpha Cen B, is very bright (Mv =1.3). Because of its brightness the frame reaches almost saturation in 15 s. The high flux level on the object fiber was observed to produce contamination of the simultaneous thorium on fiber B, which in turn introduced errors of the order of 1 to 2 m s-1 on the drift measurement. Therefore, this target is observed as ‘OBJA’, i.e. without simultaneous ThAr on fiber B.

The results in terms of radial velocity are shown in Figure 23 and Figure 24, respectively. It shall be noted that the long-term variation on HD 85512 are due to the low-mass planetary companion. Nevertheless, the some ‘high-frequency’ variation are observed on the measurements using the ThAr which are not seen on the FPCS measurements. For alpha Cen B it shall be noted that the long-term variations are due to the fact that this star is part of a triple system. The short-term variations, the scatter during the night, are however due to stellar pulsation while the night-to-night variations are due to stellar jitter.

The most important result is that in both cases neither a systematic offset nor a higher scatter is observed when using HARPS in the simultaneous FPC mode.

Figure 32: Observations of the star HD 85512 over a ten-days period with the simultaneous reference technique using the ThAr (red) and the FPC (blue). Radial-velocities are identical, although the dispersion seems to be even a little bit lower when using the FPC.

Figure 33: Observations of the star HD 128621 (alpha Cen B) over a ten-days period with the simultaneous reference technique using the FPC (blue) and without any simultaneous reference at all. radial-velocities are identical, indicating that the FPC work well even on bright objects and that it does not introduce any systematic offset or error.

# Compliance matrix

In the following we will give a compliance matrix for the FPC requirements. Most of them have already been verified on the FPC manufactured for HARPS and HARPS-N. The extension in wavelength for ESPRESSO is easily achievable by specifying the Fabry-Pérot coating for he larger wavelength range. The necessary flux will be naturally provided by the LDLS lamp, which shows almost flat distribution and no spectral features in the full ESPRESSO wavelength range.

Table 2: Compliance matrix

| Item | Required | Actual | Verification |  |
| --- | --- | --- | --- | --- |
| Wavelength coverage | 380 – 780 nm | 380 - 780nm | Test : 380 - 690 nm HARPS(-N) Design :380 - 780 nm | C |
| Photon noise | < 5 cm/s | 5 cm/s (2 cm/s) | Test : HARPS (HARPS-N) | C |
| Dynamic range PV | 4 | 6 | Test :HARPS-N | NC |
| Line width | <2/3 of spectral element | 2/3 of spectral element | Test : HARPS(-N) | C |
| Line separation | 2-5 x spectral element | 2-s5 x spectral element | Test : HARPS(-N) | C |
| Short-term repeatability | 5 cm/s per night | 3.2 cm/s per night | Test : HARPS | C |
| Long-term repeatability | 10 cm/s | 50 m/s | Test : HARPS | NC |
| Local wavelength accuracy | 10 m/s | < 10 m/s | Test : HARPS | C |

Table 4: Requirements towards the FPC

| Item | Requirement | Measure/designed |  |
| --- | --- | --- | --- |
| Wavelength coverage | 380 – 780 nm | By design (HARPS-N: 380 – 680 nm) | C |
| System transmittance | > 10% | Typical 50%, min. 10% | C |
| Transmittance uniformity | Tmax/Tmin < 2 | 5 (due to absorption of the coating in the bluest orders) | NC |
| Transmittance variations | dT/d< 2% per nm | < 0.2% per nm (see Figure 20) | C |
| Etalon total Finesse | 10 < F < 12 | By design (No problem, private communication by ICOS) | C |
| Etalon spacing | D = 7.6 mm  0.0005 mm | By design | C |
| Photon noise | < 5 cm/s | 2 cm/s (Demonstrated on HARPS-N) |  |
| Flux homogeneity | Fmax/Fmin < 5 | In the range of 390 – 680 nm | PC |
| Line width | < 2/3 of spectral element | By design | C |
| Line separation | 2-7 x spectral element | By design | C |
| Short-term repeatability | d/dt < 2 10-10 7 m/s | 0.0032 m/s in 12 hours | C |
| Long-term repeatability | d/dt < 2 10-10 7 m/s | Same as short-term, after daily calibration with a reference | NC |
| Line shape stability | Finesse determined by FR | By design | C |
| Local wavelength accuracy | λ - λ0 3 x -8λ (10 m/s) | About 10 m/s, after calibration | PC |

Three non-compliances with respect to these internal requirements must be remarked at this stage:

  - The dynamic range, i.e. the ratio between maximum and minimum flux on the CCD of the FPC is much more than 5, in particular if considering the blue wavelength range. This can however be adjusted with better adapted balancing filters. An effort shall be made with this respected while noting that a slightly higher dynamic range than specified would not have any impact of other performances.

  - The long-term stability of the FPC cannot be guaranteed, given the fact that it is a fully passive system with no feed-back from an absolute source. Coupling of the FPC to an absolute source must therefore be provided. A possible absolute source would be an atomic-clock-locked and stabilized laser. It must be recalled, however, that even a ThAr lamp may deliver the required long-term accuracy. Figure 1 shows a 10-years series on Tau Ceti. The overall dispersion is of 1.01 m s-1 and no systematic slope is measured. In fact, a linear fit delivers a slope value of 0.002 m s-1 ± 0.02 m s-1 per year.

  - The local wavelength accuracy can only be provided after calibration of the FPC with and external source or with a high-precision FTS. Both exist. The important point is, fortunately, that the dispersion law of the FP remains constant with time, such that a calibration remains valid. Only the offset would need periodical calibration.

Figure 34: 10 years of Tau Ceti demonstrating the long-term stability of the ThAr calibration.

# Conclusion

The Fabry-Pérot has demonstrated to be a valuable fall-back solution for the LFC under two aspects:

    - It is able to provide a drift measurement (simultaneous reference) well within the required photonic precision and short-term stability.

    - If combined with external absolute reference, it can be used to characterize the CCD and provide wavelength calibration. The global accuracy will be given by the external source. The local accuracy is given by the ability of calibration the group-delay dispersion of the Fabry-Pérot. Test on HARPS show that the dispersion law is reproduceable within 1/30’000’000 of the Fabry-Pérot thickness, which corresponds to an equivalent local repeatability of 10 m s-1.

# Appendix

## Data sheets

In the following we will present the data sheets of:

- Energetiq EQ99 Plus lamp controller

- Lakeshore 335 Temperature Controller

- DPG109 Pressure Controller

- MPT100 Pressure Gauge