Centro de Astrofísica da Universidade do Porto

Universidade de Lisboa, SIM/IDL & LOLS

INAF, Osservatorio Astronomico di Trieste

INAF, Osservatorio Astronomico di Brera

Observatory of the University of Geneva

Physics Institute, University of Bern

Instituto de Astrofísica de Canarias

European Southern Observatory

ESPRESSO

# Fabry-Perot Calibration

# Final Design Description and Performances Analysis

# VLT-TRE-ESP-13520-xxx, Issue 0

# March 14th, 2013

Name				Date				Signature

Prepared	F. Pepe	14/03/2013

Name				Date				Signature

Approved	F. Zerbi

Name				Date				Signature

Released	F. Pepe

# Change Record

| Issue/Rev. | Date | Section/Page affected | Reason/Remarks |
| --- | --- | --- | --- |
| 0.1 | 14/03/2013 | All | First draft version |

# Table of Contents

Chapter 1.	Introduction	7

1.1 Scope of the Document	7

1.2 Documents	7

1.2.1	Applicable Documents	7

1.2.2	Reference Documents	7

1.3 Acronyms and Abbreviations	7

1.3.1	Acronyms	7

Chapter 2.	General Requirements	10

2.1 General characteristics of the Fabry-Pérot Calibrator	10

2.2 Dimensioning of the Fabry-Pérot etalon	11

Chapter 3.	Technical Requirements	13

3.1 Operational conditions	13

3.1.1	Atmospheric pressure	13

3.1.2	Temperature	13

3.1.3	Relative humidity	13

3.2 Performance requirements	13

3.2.1	Optical requirements	13

3.2.2	Fabry-Perot parameters	13

3.2.3	Wavelengths stability	13

3.2.4	Line shape and line shape stability	14

3.2.5	Vacuum	14

3.2.6	Flux	14

3.2.7	Monitoring	14

Chapter 4.	Design architecture	15

4.1 Product Tree of the RV-reference module	15

4.2 Primary source and spectral flux	16

4.3 Etalon parameters	17

4.4 Thermal control	18

4.5 Pressure sensor and controller	18

4.6 Summary of hardware devices	19

Chapter 5.	Mechanical design	20

5.1 Overview	20

5.2 The vacuum tank	23

5.3 Optical feed-through	23

5.4 Fabry-Perot assembly	23

5.5 The entrance fiber-holder spider	24

5.6 The parabolic mirror mounting	24

5.7 The etalon mounting	25

5.8 The exit fiber-holder spider	26

Chapter 6.	Interfaces	27

6.1 Optical interface	27

6.2 Mecanical interface	27

6.3 Control system software interface	27

6.4 Hardware interfaces	27

Chapter 7.	Performance Analysis	28

Chapter 8.	Compliance matrix	29

Chapter 9.	Appendix	30

# List of Figures

Figure 1: Functional diagram of the RV reference unit	15

Figure 2: EQ99 light source and lamp controller	16

Figure 3: Spectral power distribution of the LDLS source.	17

Figure 4: Internal transmission of the FPC fibers	17

Figure 5: Internal transmission of Infrasil Fused Silica	17

Figure 6: Lakeshore 335 Temperature Controller	18

Figure 7: MPT vacuum gauge (right) and its DPG109 controller (left)	18

Figure 8: FPC hardware devices	19

Figure 9: Section view of the Fabry-Perot	22

Figure 10: Ouside view of the vacuum vessel	23

Figure 11: Fiber-optical vacuum feed-though	23

Figure 12: General view of the FPC opto-mechanical assembly	24

Figure 13: View of the upper spider and fiber support	24

Figure 14: Section view of the parabolic mirror and its support	25

Figure 15: Section view of the etalon and its support	25

Figure 16: Section view of the upper spider and fiber support	26

# Introduction

## Scope of the Document

To be integrated elsewhere …

## Documents

The applicable and reference documents are listed below:

### Applicable Documents

| AD-1 | ESPRESSO Statement of Work | VLT-SOW-ESO-13520-5059 | 1 | 01.02.2011 |
| --- | --- | --- | --- | --- |
| AD-2 | ESPRESSO Technical Specifications | VLT-SPE-ESO-13520-4633 | 3 | 01.02.2011 |

### Reference Documents

| RD-1 |  |  |  |  |
| --- | --- | --- | --- | --- |

## Acronyms and Abbreviations

### Acronyms

AD	Applicable Document

ADP	Acceptance data package

AIV	Assembly, Integration and Verification

AT	Acceptance Test

BDR	Baseline Design Review

CCB	Configuration control board

CI	Configuration Item

CIDL	Configuration Items Data List

CIL	Configuration Item List

CRB	Change Request Board

CRE	Change Request

DCL	Declared Component List

DIN	Deutsche Industrie Norm

DMPPL	Declared Materials, Parts and Process List

DRB	Delivery Review Board

ECSS	European Cooperation for Space Standardisation

EEE	Electrical electronic electromechanical components

E-ELT	European Extremely Large Telescope

EN	Euro Norm

ESO	European Southern Observatory

ESPRESSO	Echelle Spectrograph for Rocky Exoplanets and Stable Spectroscopic Observations

FAC	Final Acceptance Chile

FDR	Final Design Review

FMEA	Failure Mode and Effect Analyses

FMECA	Failure Modes Effects and Criticality Analysis

FTA	Failure Tree Analyses

HW	Hardware

I/F	Interface

ICD	Interface Control Document

ISO	International Organisation for Standardisation

IST	Integrated System Test

IV	Integration and Verification

MAIT	Manufacturing, Assembly, Integration, Test

MAIV	Manufacturing, Assembly, Integration and Verification

MAV	Manufacturing, Assembly and Verification

MMPP	Materials, Mechanical Parts and Processes

MoM	Minutes of Meeting

MRB	Material Review Board

MTBF	Mean Time Between Failures

MTTF	Mean Time To Failure

MTTR	Mean Time To Repair

N/A	Not Applicable

NCR	Non Conformance Report

NRB	Non Conformance Review Board

PAC	Provisional Acceptance Chile

PAE	Provisional Acceptance Europe

PAV	Procurement, Assembly and Verification

PDR	Preliminary Design Review

PI	Principal Investigator

PM	Project Manager

QA	Quality Assurance

RAMS	Reliability Availability Maintainability Safety

RD	Reference Document

RfW	Request for Waiver

S/S	Subsystem

SOW	Statement Of Work

SW	Software

TBC	To Be Confirmed

TBD	To Be Defined/To Be Developed

UT	Unit Telescope (8.2 meter telescope at Paranal)

VLT	Very Large Telescope

VM	Verification Matrix

WBS	Work Breakdown Structure

WP	Work Package

# General Requirements

## General characteristics of the Fabry-Pérot Calibrator

The FPC is designed as a fall-back solution in case of an LFC of suitable properties would not be available by the end of the ESPRESSO project or beginning of operations. The purpose of the Fabry-Pérot Calibrator FPC is to provide a large number of ultra-stable lines in the wavelength range of the ESPRESSO spectrograph. The FPC shall ensure ‘at least’ adequate short-term stability, i.e. better than 10 cm/s during an observing night. Since the FPC system is fully passive, no guarantee can be given a priori that the stability of 10 cm/s is achieved over long term or even years. It will therefore be operated in addition to an absolute reference, so a ThAr lamp or a stabilized laser to which it will be referenced periodically.

We remind here the properties desired for this module:

- The calibration source must cover the full wavelength domain of the spectrograph.

- The line position and shape stability of all lines must be better than 0.1 m/s over the instrument life-time.

- The lines are not resolved by the spectrograph, since otherwise information is lost, respectively we do not take full advantage of the spectral resolution of the instrument. The resolution of the spectrograph is 120’000 in its standard mode, limited by the illuminating fiber size.

- The line distance is perfectly stable and analytically defined.

- The line distance must be minimum 2 and maximum 4 FWHMs of the spectrograph IP.

- The relative intensity of any neighboring lines must be stable at 10% over time.

- The dynamic range of line intensities over one order must be smaller than a factor 2.

As a first goal, the system shall be made sufficiently stable to be used as a relative reference during an astronomical observing night of typically 12 hours. The module will possibly need to be re-calibrated every night, e.g. with a ThAr lamp or a LFC, for absolute referencing, but will serve as perfect simultaneous reference within the night. The second objective will be to extend the time during which the module remains stable. Ideally, the module will be intrinsically so stable, that re-calibration will not be needed over large time scales (years).

Reflection is obtained by a multi-layer thin film coating of the inner optical surfaces. The following aspect must be considered: The apparent separation of the two spacers depends on the effective penetration depths of the wave on the reflective surface.

- Because of the varying real and imaginary part of the refractive index as a function of wavelength, the effective FP separation will vary with wavelength. This wavelength dependence must be avoided as far as possible, in particular, a temperature or aging effect must be avoided. If this wavelength dependence cannot be avoided completely, it should be at least very “smooth”.

- Thermal effects in the coating will affect the way the etalon behaves: The dielectric stack will have a certain thickness. When the ambient temperature changes, the stack will expand (CTE of some 10-6) and this will change the effective gap of the etalon. With a gap of 7 mm, for instance, and a dielectric stack of 4 µm thick with a CTE of 7 10-6, the variation of the apparent separation will be about 8 10-9 / K. In the same way, the index of refraction will change with temperature and affect the apparent penetration depth of the optical wave in the coating.

- Aging of Zerodur has been studied in metrology labs. Schott data indicates shrinkage of 10-7 per year. However, the experience on the HARPS FP shows that the effect of aging is well below the estimated value. In alternative to Zerodur, ULE shall be used.

When making trade-off, the following order of priorities shall be considered:

- Relative wavelength stability over temperature and time (between two neighboring lines)

- Wavelength stability over temperature and time

- Equidistance of transmitted wavelengths

- Uniformity of line intensity and line shape and width

## Dimensioning of the Fabry-Pérot etalon

The primary parameters for which the Fabry-Pérot etalon has to be dimensioned are :

Variables:

- Spectrograph resolution RS

- Shortest wavelength λB

- Longest wavelength λR

Requirements for an optimum use of the lines are:

- The line separation is minimum on the blue side, because of the FP equation. The separation must be chosen such that the lines can be resolved by the spectrograph:
The peak separation at all wavelength must be > 3 FWHM: Δλ≥3∙λRS 	(R1)

- The line separation is highest on the red side. At a given Finesse, the FP lines become wider there. It shall be avoided, that the spectrograph resolves the lines:
The FP line must be under-resolved by a factor of 2/3 : δλ≤2/3∙λRS	(R2)

Let us now first derive the required FP. Given the gap D of the etalon, the resonance of the FP will be achieved whenever	2D=mλ 	(1).

Therefore, the separation between the 2 bluest peaks will be, using requirement (R2):

Δλ=2Dm-1-2Dm=2Dmm-1≥3∙λRS 	(2)

Let's replace m in the last inequality by using equation (1),

2Dmm-1≅2Dm2=λ2/2D≥3∙λRS	 (3)

and solve for D. In order to be true for all wavelengths, the relation must me satisfied for the shortest wavelength λ=λB, such that we get   D=λB∙RS/6 	(P1)

Now let's derive the required finesse F of the etalon: The Finesse is defined as the ratio of the line separation to the FWHM:	F=Δλ/δλ	 (4)

We can express again Δλ as a function of the gap and the FP order

Δλ=2D/m(m-1)≅2D/m2 	(5)

and introduce equation (5) and requirement (R2) into equation (4)

F=Δλδλ≥2D3RSm22λ=3DRSm2λ=3DRSλ24D2λ=3RSλ4D 	(6)

where in the step before the last we used the etalon equation (1). Since this condition must apply at all wavelengths, we have to choose F for the longest wavelength and substitute D using (P1):

F=3RSλR4D=18∙λR4λB=9λR2λB	(P2)

Let's now look at the etalon parameters from the instrument figures: RS = 120'000, λB=380nm, λR=780nm we find

D = 7.6 mm

F = 9.2

# Technical Requirements

## Operational conditions

The FPC will be operated in the CCL at the Paranal Observatory, Chile. It will be located within 2T. The FPC will not be exposed to any motion, gravity, vibrations, etc.

### Atmospheric pressure

Atmospheric pressure is 760 mbar in average. Given the high ambient pressure variations the FPC etalon will be operated in vacuum. A pressure value of 0 mBar must be considered for the optical computations.

### Temperature

The temperature of the environment is typically of TBD ± TBD°C.

### Relative humidity

The system is working normally in a controlled environment with relative humidity always lower than 100%. During manufacturing, transportation, installation, and maintenance the humidity can attain peeks close to 100%.

## Performance requirements

### Optical requirements

Operational wavelength range:	380 nm - 780 nm

Total transmittance at peak:	T > 10% 	in the wavelength range

Transmittance uniformity	Tmax/Tmin < 2	in the wavelength range,

may be balanced with

the light source or filter

Transmittance variations	dT/d< 2% per nm	in order not to distort the

Lines by more than 10 cm/s

### Fabry-Perot parameters

Effective total Finesse:	10 < F < 12	 in the wavelength range

Fabry-Perot spacing:	D =  7.6 mm  0.0005 mm

In order to achieve a total (reflectivity-limited) Finesse of F = 10, the various Finesse contributions must be:

Reflectivity Finesse: 	FR≥54∙10=12.5

Aperture Finesse:	Fθ≥2.35∙F=23.5

Defect Finesse: 	FD≥2∙2.35∙F=33.2

Parallelism Finesse:	FP≥2∙2.35∙F=33.2

### Wavelengths stability

The Fabry-Perot transmitted wavelengths must remain stable with time. The Fabry-Perot shall be placed in vacuum, and the required temperature stability specified. Changes of the transmitted wavelengths must be minimized.

Wavelength repeatability:	d/dt < 2 10-10 7 m/s	at any wavelength

required: during at least 12 hours

goal: during 10 years

Absolute wavelength:	λ - λ0 	3 x -7λ (100 m/s)	at any wavelength

The absolute wavelength of a line is defined on the centroid of the transmitted FP line with respect to the expected wavelength given by the nominal etalon spacing. Wavelength repeatability indicates the stability of the centroid of a line.

The Fabry-Perot must be optimized for minimum sensitivity of the transmitted wavelengths with regard to temperature variations. If the final design is intrinsically stable it relaxes the contraints on the thermal control stability.

### Line shape and line shape stability

The shape of the transmitted FP lines (intensity, symmetry, width) shall be as uniform as possible. The line shape shall be as stable as possible. It is therefore suggested that the etalon finesse is determined by the reflectivity finesse, which is the most insensitive to geometrical aspects. However, in order to make the transmitted lines 'equidistant' to a level described by the 'absolute wavelength' requirement, the phase of the etalon transmittance function, and thus the reflectance of the individual mirrors, shall be kept constant as a function of wavelength, if possible. Since this is very difficult, the plan will be to calibrate the phase variations with an FTS at the beginning of the lifetime of the FPC.

### Vacuum

From the stability requirements we deduce that the air pressure around the FP etalon must be kept stable within 10-3 mbar during 12 hours. The requirement shall be therefore formulated as:

dp/dt < 2 x 10-3 mbar per day.

### Flux

The FPC shall deliver sufficient flux at peak wavelength within the étendue accepted by the spectropgraph, such that a global photon-noise precision of better than 5 cm/s is achieved in less than 20 s.

### Monitoring

The temperature of the FPC, and in particular the etalon, must be monitored and controlled continuously with a resolution of 0.001 K, a short-term repeatability of 0.002 K over time scales of 1 months and of 0.025 K over five years, and an absolute accuracy of 0.1 K. The monitoring must have a data link to the ESPRESSO control system.

The air pressure around the etalon must be monitored continuously with a resolution and precision of 0.0001 mbar, and an absolute accuracy of 0.001 mbar. The monitoring must have a data link to the ESPRESSO control software.

# Design architecture

The proposed FPC is derived from a new concept that has been successfully evaluated on HARPS and HARPS-N: The stability reference for the radial velocity is essentially a Fabry-Perot etalon housed in a temperature controlled vacuum enclosure. The etalon is fibre-fed by a bright white lamp. A symmetrical set-up of parabolas couples the input to the exit fibers, with the etalon located in the collimated beam between the two parabolas, making the design achromatic.

Figure 1: Functional diagram of the RV reference unit

## Product Tree of the RV-reference module

- White source

- LDLS source with SMA or FC fiber connection

- Balancing filter (color correction)

- 1-m fiber 230 µm SMA/FC-FC with metallic protection jacket

- Optics

- 1 fiber vacuum feedthrough

- 1 thin input fiber FC-FC of 230 µm dia. without jacket

- 2 protected-silver coated parabolic mirrors with f = 100 mm

- 1 etalon

- 1 thick output fiber FC-FC of 600 µm without jacket

- FP-etalon assembly (structure according to drawing)

- Vacuum tank

- Electroplated aluminum body

- Dismountable cover

- DN-40 pumping flange

- 2 fiber-feedthrough mounting flanges

- 1 electrical feedthrough mounting flange

- Temperature control device

- 1 Lakeshore 335 controller

- 1 connect cable for heater and sensors

- 2 wired Si-diodes temperature sensors

- 2 wired heating foils

- Vacuum monitoring device.

- Digiline pressure sensor

- Connection cable

- DPG-109 pressure-gauge controller

## Primary source and spectral flux

For irradiance reasons, spectral flatness and MTBF, we would like to use a innovative source called “laser driven light source” EQ-99FC Plus from the company Energetiq. This source us presently being used in integration on HARPS North with convincing results. The light source possesses a complete remote control system for control via software. However, the lamp will be used in stand-alone mode, i.e. it will manually switched ON and kept ON continuously for FPC stability reasons. The lifetime of such a lamp is longer than 20’000 hours. A LRU spare shall be procured for immediate exchange in case of failure. For the fibers standard Ceramoptec octagonal fibers with high internal transmittance will be chosen.

Figure 2: EQ99 light source and lamp controller

Figure 3: Spectral power distribution of the LDLS source.

Figure 4: Internal transmission of the FPC fibers

In order to reduce the heat-up of the FP etalon by the source, we will introduce in the laser source a heat-blocking filter TKG-5253 by UQG Optics with the transmittance curve shown is Figure 5. As a function of the obtained chromatic distribution, a balancing filter will be added. For the moment no choice has been made, since a) a balancing filter may not be necessary, given the high blue flux of the Fabry-Pérot, and b) we prefer in any case to see the resulting chromatic flux distribution. A filter can be inserted easily at anytime an a appropriate location will be reserved for it.

Figure 5: Heat-blocking filter TKG – 5253 transmittance

## Etalon parameters

Given the gap and Finesse requirement, the etalon will be designed as follows:

Gap: D = 7.6 mm

Clear diameter D = 40 mm (design), focal length of collimator f = 100 mm

Reflectivity Finesse: 	FR = 12.5	R() = const = 78%

Aperture Finesse:	F = 23	Fiber diameter d < 300 microns @ fcoll = 100 mm

Defect Finesse: 	FD = 33	Mirror wavefront better than l/50 P-V

Parallelism Finesse:	FP = 33	Departure from parallelism < 10 nm -> 0.25 10-7 rad

The etalon plates will be manufactured form fused silica blanks, most likely INFRASIL. The internal fibers being very short, we propose to use Ceramoptec WF fibers as well.

Figure 6: Internal transmission of Infrasil Fused Silica

## Thermal control

From our experience with previous similar systems, the thermal control of the etalon structure to a couple of mK rms should be a straightforward affair. It will be based on a Lakeshore 335 temperature controller using silicon diodes as sensors and one or two thermal pads as heating elements.

Figure 7: Lakeshore 335 Temperature Controller

## Pressure sensor and controller

The pressure of the FPC and thus around the etalon will be measured by means of an MPT100 Pfeiffer gauge attached to the vacuum tank. A DPG109 controller, the one used for the Vacuum System, will read out the sensor. The DPG controller will be read out through the ESPRESSO ICS.

Figure 8: MPT vacuum gauge (right) and its DPG109 controller (left)

## Summary of hardware devices

Figure 9: FPC hardware devices

| Type | Manufacture | Reference | Interface | Notes |
| --- | --- | --- | --- | --- |
| Temperature controller | LAKESHORE | 335 | USB or Ethernet |  |
| Temperature sensor | LAKESHORE | Si diodes DT-670 |  | Silicon diode |
| Pressure controller | PFIEFFER | DPG 109 | RS-232 (DB9) | 9 inputs |
| Pressure gauge | PFIEFFER | MPT100 |  |  |
| Lamp controller EQ99 Plus | Energetix |  |  |  |

# Mechanical design

## Overview

The vacuum tank is a simple cylinder of electro-polished aluminum with 2 aluminum covers. It contains the actual FPC and its etalon, which is sketched in the following. While the first model of the FPC had a structure made of optically contacted Zerodur cylinders, a sensitivity analysis demonstrated later that this was not necessary. The mechanical structure of the second model (Harps North) is more classical and a less expensive assembly of aluminum elements holding the mirrors, the etalon and the fiber connectors.

Parabolic mirror

Adjustable spider in X,Y,Z

Output fibre

Etalon

Input fibre

Parabolic mirror

Adjustable spider in Z

Aluminium body

Vacuum vessel

Input feed through

Output feed through

Figure 10: Section view of the Fabry-Perot

Description of the light path

The light coming from the front-end enters the system by a vacuum feed-through (1), then goes through an optical fibre (2), fixed at one end to a spider (3). The light coming out of the fibre is projected against a parabolic mirror (4) which sends back the collimated beam trough the etalon (5) and against the second parabolic mirror (6). This mirror focuses again the beam and the light is finally collected by the output fibre and goes out trough the second vacuum feed-through.

4

6

5

3

2

1

Light output

Light input

## The vacuum tank

The vacuum vessel is a hermetic chamber in which the pressure is maintained at 1E-3 mbar and the temperature is controlled. It is made out of Aluminum and is composed by 3 parts:

- A centre body (1): it is a tube with 2 welded flanges at each end.

- A blank flange (2) on the bottom side.

- A flange with feed-through (3) on the upper side. 2 optical feed-through (4) and 1 electrical feed-through (5). There is also a DN 40 ISO-K half nipple flange welded on for the pumping connection.

2 seals insure the airtight and 8 clamps maintain all parts together.

Figure 11: Ouside view of the vacuum vessel

4

2

5

3

1

## Optical feed-through

It is an off-the-shelf part composed by a flange and 2 FC receptacles at both ends. An optical fibre passes through a soft sealant material, which is compressed by the compressive end-cap and internal squeezer of the feed-through.

Figure 12: Fiber-optical vacuum feed-though

## Fabry-Perot assembly

It is a stand-alone system placed in the vacuum vessel and fixed to the top flange in 3 points. Some insulated housings avoids thermal transmission from the outside.

Figure 13: General view of the FPC opto-mechanical assembly

Vacuum vessel top flange

First spider

Input optical fibre

Second spider

Parabolic mirror

Etalon

Output optical fibre

Parabolic mirror

Insulated housings

## The entrance fiber-holder spider

This is the support for the input optical fibre connector (1). This spider (2) is adjustable only in vertical (Z) direction by shimming spacers (2).

Figure 14: View of the upper spider and fiber support

3

2

1

## The parabolic mirror mounting

These are the two mirrors (1) at each end of the system. The first one collects the beam coming from the input fibre and collimates it back through the etalon. The second one collect the light coming from the etalon and focus it back to the output fibre. These mirrors have a reference surface (2) on the same side as the parabola (3). A Viton® O-ring (4) compressed by an aluminium ring (5) maintain the mirror in position. 3 spacers (6) avoid a contact between the mirror (1) and the aluminium ring (5).

Figure 15: Section view of the parabolic mirror and its support

6

3

2

5

4

1

## The etalon mounting

The etalon (1) is the main optical part. At its base, it has a flange to fix it. 3 brackets (2) and a Viton® O-ring (3), used as a sprig, maintain the etalon in position on the supporting flange (4).

Figure 16: Section view of the etalon and its support

4

2

3

1

## The exit fiber-holder spider

This is the support for the output optical fibre connector (1). This spider (2) is used for the final optical adjustment. It is adjustable in all directions (X,Y,Z): 4 pushing screws (3) for X,Y direction and 3 spacers (4) to shim for the vertical direction. The 4 screws are removed after adjustment. To insure an accurate displacement in X and Y direction, an intermediate ring (5) is placed between the supporting flange and the spider. In this intermediate ring are machined grooves (6) in the 2 directions.

Figure 17: Section view of the upper spider and fiber support

6

4

3

2

1

5

# Interfaces

## Optical interface

The FPC will have only one optical interface to the calibration unit. It will provide a male FC connector holding a 600 µm fiber to be connected at the input port of the calibration unit.

## Mechanical interface

The opto-mechanical part of the FPC will be housed in a 0.6m x 0.6m x 1m enclosure of 60 x 60 x 80 cm. This enclosure is mainly filled with insulating foam and hosts the vacuum tank in the middle. This enclosure weights about 25kg and must be located in a stable thermal environnment (2T).

The primary light source will be packaged in a 6U 19” rack. The lamp control system will be packaged in a 3U 19” crate, as well as the Lakeshore temperature controller and the vacuum gauge controller

The Fabry-Pérot will be attached to the pumping system of the vacuum system for initial pumping and possible periodic pumping. Normally, no pumping will be performed on a regular basis, however, given the fact that the pressure increase is sufficiently small to avoid nightly drifts larger than the maximum requird10 cm/s, and, becse long-term drifts will be calibrated with an external absolute calibration reference

## Control system software interface

No active control is to be provided. The FPC will only be connected ‘passively’ to the ICS for monitoring of:

- Lakeshore 335 temperature controller for the two temperatures of the FPC

- Pressure DPG109 for the readout of the pressure gauge. The pressure gauge will be actually be connected to the Vacuum System pressure controller.

## Hardware interfaces

The FPC is a stand-alone and passive unit. No remote power control is required since it will be continuously operated. The unit is turned ON and turned OFF manually, whenever necessary. The controllers are installed in standard 19” rack within a small cabinet. A single standard 230V50Hz UPS power line must be delivered to the cabinet

The Lakeshore 335 is accessible via an Ethernet or serial port.

The Pfeiffer DPG109 pressure controller of the Vacuum System will be used to read-out also the FPC pressure gauge. The DPG is controller through a serial port.

# Performance Analysis

In the following we shall present the results obtained with a FPC already produced for HARPS and HARPS-N. These results replace any analysis, since it provides real-life and assessed performances.

Figure 18: FPC currently in use on HARPS-N during the tests (left) and in its isolation box (right)

## Flux measurements

The flux per extracted pixel at the peak of the echelle order is shown for a 40 s exposure with the FP system in the present configuration. A RV precision of about 4 cm s-1 per frame is obtained on a single frame, which is about twice a good as using the ThAr for the simultaneous drift measurement. On the other hand, the flux results to be about a factor of 10 lower than with the Xe-lamp, which has an larger etendue which fills in a better way the calibration fiber.

By optimizing the setup a value below 2 cm s-1 can be attained (tested with Xe-lamp). In particular, the blue flux should be improved by increasing the source power and by using a blue-balancing filter. In the present situation, the orders below number 10 are not used for the drift computation.

Figure 19: Peak flux per extracted pixel as a function of the order. The results were obtained using the SC laser source and an exposure time of 40 s.

Figure 20: FPCS flux per exctrated pixel along the main dispersion at the blaze peak for echelle order #50

## Stability measurements

### Drift computation and use of new algorithm

In order to use the FPCS in a standard way on HARPS, the DRS (pipeline) had to be adapted and some functionality added. In particular, one had to add the possibility of treating frames using the FPCS on fiber B for both calibration and scientific (target) exposures. The recognition and the correct handling of the various cases had to be implemented. Also, for the measurement of the drift, several new algorithms have been tested and implemented.

It is important to note that these development have been also used for the simultaneous development of the LFC by ESO and MPQ, since the LFC as well had to be tested on HARPS.

During the test phases we remarked that all the used algorithms for measuring the spectrograph drifts on fiber B would depend on the ‘spectral content’ of the reference spectrum (LFC, ThAr, FPCS)  and sensitive to the flux. We have therefore developed a new algorithm based on the Fourier transform (Chelli, A&A, 2001) which can easily be applied to any kind of spectrum and shows better performance and robustness than any other algorithm. The algorithm is presented in the Appendix of the present document.

### Flux-dependence tests

Triggered by tests made with the LFC, which showed some dependence of the drift-measurement on the flux, we have performed tests using the FPCS during which the exposure time (and thus the integrated flux) was varied by a factor of 100. Several algorithms showed a strong dependence on flux to up 10 m s-1 over a the applied flux range. With the new algorithm, however, this effect was measured to be of only 2-3 m s-1 in total. Figure 18 shows the results, where the x-axis indicates the relative flux level (or exposure time) and the y-axis the drift error expressed in the unit of m s-1.

Figure 21: Flux dependence of the drift measurement determined using the FPCS

It is still not clear which is the origin of this effect. Possibilities are a flux-dependent CTE or effects of varying background in drift-measuring algorithm. Nevertheless, the effect is small, especially if we consider that in normal operation the flux of the reference spectrum varies less than 10% and that of the star typically by a factor 2, in which case the error would be of the order of 0.5 m s-1.

### Short-term stability

We have tested the short-term stability by performing long series of ThAr-FP exposures (ThAr on fiber A, FPCS on fiber B) and by computing the respective drifts independently. Figure 19 shows the drift from one exposure to the other as computed by the ThAr lamp and the FPCS, as well as the difference of both to remove possible instrumental drifts.

For the ThAr alone we obtain a dispersion of about 0.19 m s-1 rms which should be compared to about 0.09 m s-1 pure photon-noise error. For the FPCS we obtain 0.176 m s-1 rms dispersion, to be compared to 0.06 m s-1 photon noise. For the difference of both, we end up with a photon noise of 0.11 m s-1 and a measured dispersion of 0.14 m s-1. The dispersion of the dirrence is well below the square root of the quadratic sum of the individual drift measurements of 0.26 m s-1, demonstrating that both sources track a ‘real’ instrumental’ drift. On the other hand, the difference between the photon noise and the measured dispersion leaves the room for a 0.08 m s-1 additional dispersion, which is not ‘seen’ by at least one of the fibers (or better sources). We suspect here a residual drift of the FPCS or the ThAr sources (heating, instabilities, etc.) or very small instabilities of the light injection which do not affect both sources or fibers in the same way.

Figure 22: Relative drift measurements over 7 hours to test the short-term performances of the FPCS

The plot in Figure 20 shows the same data, but this time in absolute terms, i.e. it indicates the absolute drift for FP and TH with respect to the very first exposure of the series. It should be remarked that in general the FP and the TH follow the same ‘high-frequency’ behavior, tracking real instrumental drifts. This is demonstrated by the fact that the dispersion of the difference is well below the dispersion of the individual measurements. On the other hand, the plot shows also that one can have differential drifts of the order of 0.9 m s-1 P-V or 0.23 m s-1 rms if no recalibration is performed during the night. Again, this is a worst case limit if we assume that the differential drift between TH and FP is fully due to the FP, we is very unlikely.

Figure 23: Absolute drift measurements over 7 hours to test the short-term performances of the FPCS

### Long-term stability

The long-term stability was tested in a similar way as the short-term stability. This time, however, the drift was computed with respect to a reference frame at the beginning of the two-month period. Frames were taken every 5 days. The idea was to monitor possible long-term drifts of the FPCS with respect to the ThAr2 lamp.

Figure 21 shows the results. Over the almost 3 months from March 24 to June 15, 2011, both ThAr and FPCS drifted by about 4 m s-1, reproducing most likely a real drift of the spectrograph. Relative to each other, the drift was smaller that 2 m s-1 at the end of the period. This latter value shows clearly that the long-term drift of the FPCS cannot be larger than typically 3 m s-1 over 60 days, and must be thus well below the specified 0.1 m s-1 per night.

On the other hand we observe that the differential drift has attained values up to 2 m/s. It is not clear where this comes from. Again, we suspect small instabilities of the light injection which do not affect both sources or fibers in the same way, but we cannot exclude effects of any of the sources (FPCS or warming-up effects of the ThAr lamp). It should be noted, however, that the differential drift is monitored with respect the same reference frame, and that therefore, we do not have to expect differential drift measurement errors of the same magnitude on a night-to-night basis, for which we have measured values of 0.14 m s-1 rms.

Finally, we would like to note that the scatter of the ThAr measurement seems to be higher than that of the FP. We suspect lamp ‘instabilities’ to influence the ThAr precision!

Figure 24: Drift measurements over 80 days to test the long-term performances of the FPCS

### Effective etalon gap

In the FP-7 documentation we reported about the measurement of the effective etalon gap as a function of wavelength by using the ThAr calibration of the FP frames. We have repeated this measurement in order to verify whether the etalon characteristic, in particular the coatings, remain constant with time. Figure 22 shows the effective gap (twice the physical gap) as a function of wavelength. The two curves and the respective zooms show a measurement taken with the Xe-lamp in 2010 and a measurement taken with the SC-laser in 2011 one year later.

The general offset can be explained by the differences in pressure inside the etalon dewar between the two measurements. It interesting to note that the two curves show the same low-frequency behavior, confirming that the coatings characteristics remain unchanged. The medium-frequency variations seen in the second plot are probably due to the line asymmetries introduced by the blaze function; they as well remain very stable, as expected. Finally, the origin of the high-frequency variations seen in the last plot is not understood yet, but again they seem to remain stable in wavelength and amplitude over the one year period.

Figure 25: Effective etalon gap as a function of wavelength

## Evolution of the FO chromatic dispersion with time

Bruno, graphique et analyse

# On-sky tests and operations

## Description of operations

In order to operate the FPCS in a standard way, we have implemented a operational scheme in an identical way as when the ThAr lamp is used. Templates had to be added and the DRS modified in order to do so. In the following we describe the operations step-by-step:

- During the calibrations, a wavelength-calibration exposure is taken where fiber A, the object fiber, is illuminated with the lamp ThAr 1 and fiber B, the reference fiber, with the FPCS. This exposure is produced by calling the calibration Template ‘HARPS_ech_cal_WAVE’ with the fiber B parameter set to ‘FP’ by default.

- The DRS recognizes this frame as such and performs extraction and wavelength calibration on fiber A and extraction and referencing on fiber B. The result is stored in the calibration database calib_DB under the reference ‘wave’. It shall be noted that for the wavelength calibration of fiber B the latest ‘THA2’ exposure (both fibers illuminated with ThAr) is used.

- During the scientific observation fiber A, the object fiber, is illuminated with the astronomical target and fiber B, the reference fiber, with the FPCS. This exposure is produced by calling the acquisition template ‘HARPS_ech_acq_WAVEsimult’ with the fiber B parameter set to ‘FP’ by default.

- The DRS recognizes this frame as such and performs extraction and stellar velocity computation on fiber A and extraction and drift computation on fiber B. For the latter, the reduction recipe looks into the calib_DB for the latest ‘WAVE’ calibration with ‘FP’ on fiber B and computes the differential drift between the present exposure and the reference exposure.

This scheme has been implemented and is fully operation. Two ‘HARPS_ech_cal_WAVE’  using ‘FP’ on fiber B have been included in the Standard-RV calibration set and are executed every night on HARPS.

## Observations of known standards

In order to test the on sky performance we have observed, during the ‘Sousa’ mission in March 2011, two well-known targets. Both targets are part of the ‘HARPS-Upgrade GTO’ program by Pepe et al. In order to compare the radial-velocity results, we have observed both targets in their standard modes (usually used modes). The first target, HD 85512, is a moderately bright stable star with one known planet. Usually, this target is observed in ‘thosimult’ mode. The second star, HD 128621 or alpha Cen B, is very bright (Mv =1.3). Because of its brightness the frame reaches almost saturation in 15 s. The high flux level on the object fiber was observed to produce contamination of the simultaneous thorium on fiber B, which in turn introduced errors of the order of 1 to 2 m s-1 on the drift measurement. Therefore. This target is observed as ‘OBJA’, i.e. without simultaneous ThAr on fiber B.

The results in terms of radial velocity are shown in Figure 23 and Figure 24, respectively. It shall be noted that the long-term variation on HD 85512 are due to the low-mass planetary companion. Nevertheless, the some ‘high-frequency’ variation are observed on the measurements using the ThAr which are not seen on the FPCS measurements. For alpha Cen B it shall be noted that the long-term variations are due to the fact that this star is part of a triple system. The short-term variations, the scatter during the night, are however due to stellar pulsation while the night-to-night variations are due to stellar jitter.

The most important result is that in both cases neither a systematic offset nor a higher scatter is observed when using HARPS in the simultaneous FPCS mode.

Figure 26: Observations of the star HD 85512 over a ten-days period with the simultaneous reference technique using the ThAr (red) and the FPCS (blue). Radial-velocities are identical, although the dispersion seems to be even a little bit lower when using the FPCS.

Figure 27: Observations of the star HD 128621 (alpha Cen B) over a ten-days period with the simultaneous reference technique using the FPCS (blue) and without any simultaneous reference at all. Radial-velocities are identical, indicating that the FPCS work well even on bright objects and that it does not introduce any systematic offset or error.

# Compliance matrix

| Item | Required | Designed |
| --- | --- | --- |

# Appendix

## Data sheets

In the following we will present the data sheets of:

- Energetiq EQ99 Plus lamp controller

- Lakeshore 335 Temperature Controller

- DPG109 Pressure Controller

- MPT100 Pressure Gauge