Executive Summary

Doc. nr. VLT-TRE-ESP-13520-0037

Preliminary Design

# ESPRESSO

Introduction

At its 67th Meeting in October 2007, the ESO STC recommended the development of additional second-generation VLT instruments and its detailed proposal was endorsed by the ESO Council at its 111th meeting in December 2007. Among the recommended instruments, a High Resolution, Ultra Stable Spectrograph for the VLT combined Coudé focus arose as a cornerstone to complete the current 2nd generation VLT instrument suite. Following these recommendations, ESO issued on March 2007 a call for proposal, open to member state Institutes or Consortia, to carry out the Phase A study for such an instrument. The submitted proposal was accepted by ESO and the Consortium led by the Observatoire de Genève was selected to carry out the Phase A Study for this High-Resolution Ultra Stable Spectrograph.

The main scientific drivers for this instrument have been defined by ESO as follows:

1.	Measure high precision radial velocities for search for rocky planets

2.	Measure the variation of physical constants

3.	Analyze the chemical composition of stars in nearby galaxies

The Echelle SPectrograph for Rocky Exoplanet and Stable Spectroscopic Observation, ESPRESSO, described in the Preliminary Design Report, is a fiber-fed cross-dispersed echelle spectrograph which can be operated with one or up to 4 UTs of ESO’s VLT. It will be located in the Combined-Coudé Laboratory (CCL) of the VLT to be the first instrument able of using a 16-m equivalent telescope. The instrument design and the project organization have been chosen to comply with the main scientific drivers mentioned above, but offer many new and exciting scientific possibilities to the community.

Scientific context

High-resolution spectrographs have always been at the heart of astrophysics. In particular, they form the observational basis for stellar physics and the interstellar and intergalactic medium. Correspondingly, such facilities have always been on high demand at major observatories, see e.g. UVES at the VLT or HIRES at Keck. As telescopes get larger, the capabilities of high-resolution spectrographs extend to fainter objects, e.g. faint dwarfs in the solar neighborhood, resolved stars in local galaxies and extragalactic objects. Besides the increase in photon collecting power, another aspect has emerged in recent years: the need for high-precision spectroscopy, i.e. spectra where the instrumental effects are considerably reduced providing highly repeatable observations over long timescales. One prominent technique where such considerations are particularly relevant is the measurement of high-precision radial velocities, or more generally the determination of the positions and shapes of spectral lines. The HARPS spectrograph at the ESO-3.6m telescope has been a pioneering instrument in this respect. It clearly appears today that high-precision spectroscopy has a great potential for new exciting discoveries.

The need for a ground based follow-up facility capable of high RV accuracy has been stressed in the ESO-ESA working group report on extrasolar planets. In fact, terrestrial planets in habitable zone is one of the main scientific topics of the next decades in Astronomy, and one of the main science drivers for the new generation of ELTs. The ESO-ESA document at p. 63 states: “ High precision radial velocity instrumentation for the follow-up of astrometric and transit detections, to ensure the detection of a planet by a second independent method, and to determine its true mass. For Jupiter-mass planets, existing instrumentation may be technically adequate but observing time inadequate; for Earth-mass candidates, special purpose instrumentation (like HARPS) on a large telescope would be required.” The same concept is reiterated in the first recommendation: “Support experiments to improve radial velocity mass detection limits, e.g. based on experience from HARPS, down to those imposed by stellar surface phenomena” (ESA-ESO report, p. 72).

Do the fundamental constants vary? This is one of the six big open questions in cosmology as listed in the ESA-ESO report for fundamental cosmology (Peacock et al 2006). In the executive summary, the document states: quasar spectroscopy also offers the possibility of better constraints on any time variation of dimensionless atomic parameters such as the fine-structure constant α and the proton-to-electron mass ratio. Presently there exist controversial claims of evidence for variations in α, which potentially relate to the dynamics of dark energy. It is essential to validate these claims with a wider range of targets and atomic tracers." This can be done only with improved spectroscopic capabilities.

ESPRESSO combines an unprecedented accuracy in the radial-velocities determination together with the largest photon collecting area available today. It will certainly provide breakthroughs in many areas of astronomical research, many of which we cannot anticipate.

ESO

ESO

The Context

Use of GTO

For the capital and human investment, the Consortium will be awarded Guaranteed-Time Observations. 80% of the observing nights will be invested for the search and characterization of rocky planets in the habitable zone of G, K, and M stars in the 1-UT mode. A 10% of the time will be dedicated to the determination of possible variability of fundamental constants. Depending on the magnitude of the target, this programme will be carried out partially in the 1-UT, partially in the 4-UT mode. The remaining 10% will be reserved for outstanding science cases and allocated as a function of topical questions arising at the moment of the GTO Observations.

- Chemical composition of stars in local galaxies

- Investigation of metal-poor stars

- Stellar oscillations, asteroseismology

- Diffuse stellar bands in the interstellar medium

- Chemical enrichment of IGM

- Galactic winds and tomography of the IGM

- Chemical properties of protogalaxies

- Cosmology and expansion of the Universe

- Further extra-solar planet science

- The Rossiter-McLaughlin effect

- Transmission spectroscopy

- Planets in nearby open clusters and galaxies

The search for rocky-planets in the habitable zone

Since 2004, research teams using the RV technique have been making the headlines several times for their discoveries of low-mass extrasolar planets, some of which having only a few times the mass of the Earth. Two thirds of the exoplanets known today having a mass below 10 Earth Masses have been discovered with HARPS. The rate of these discoveries increases steadily. From the high-precision program it has been shown that 50% of the stars harbor Neptune-mass planets and super-Earths. These discoveries were made possible thanks to the sub-m s-1 precision reached by this instrument. Given the faint magnitude of the star and/or the tiny radial-velocity signal induced by the planet, most of the observed objects would have remained out of reach of existing facilities that were limited to ~3m s-1. This emerging population is supposed to be only the tip of the iceberg. Considering the observational bias towards small masses, on the one hand, and the predictions made by planet-formation models, on the other hand, one should expect a huge amount of still undiscovered low-mass planets, even in already observed stellar samples.

ESPRESSO is designed to explore this new domain and to enter into unknown territory. This goal can only be obtained by combining high efficiency with high instrumental precision. ESPRESSO will be optimized to obtain best radial velocities on quiet solar-type stars. A careful selection of these stars will allow to focusing the observations on the best-suited candidates: non-active, non-rotating, quiet G to M dwarfs. With a precision of 10 cm s-1 it will be possible to detect rocky planets of few Earth masses in the habitable zone of solar type stars. By extending the sample towards the lighter M-stars the task becomes even ‘easier’. Given its efficiency, spectral resolution and spectral domain, ESPRESSO will be most efficient up to M4 dwarfs.

The discovery and the characterization of this new population of very-light planets will open the door to a better understanding of planet formation and deliver new candidates for follow-up studies by transit, astrometry, transit spectroscopy, Rossiter-McLaugling effect, etc. In addition, ESPRESSO will be the ideal (and maybe only) machine to make spectroscopic follow-up of Earth-sized planets discovered by the transit technique.

Do physical constants vary?

In 1999, observations of spectral lines in distant QSOs gave the first hints that the fine-structure constant might change its value over time, being lower in the past by about 6 parts per million (Webb et al 1999, Murphy et al. 2004). More recently, also μ= mp / me has been found to vary by Ivanchik et al (2005), Reinhold et al (2006). However, other recent results suggest no variation. The debate is still open and calls for new, more precise measurements. Only a high-resolution spectrograph that combines a large collecting area with extreme wavelength precision can provide definitive clarification.

VLT ESPRESSO, by means of its extreme stability and wavelength calibration, will provide an increase in the accuracy of the measurement of these fundamental constants by at least 1 order-of-magnitude. A confirmation of variability with high statistical significance is of crucial importance. It will open a window to new physics beyond the Standard Model and possibly shed new light on the nature of Dark Energy. More stringent bounds are important to constrain various theoretical models that lead to the variability of fundamental constants.

Some very few objects are sufficiently bright to be observed with one UT. Fainter objects will greatly benefit rom the possibility of using four UTs simultaneously and thus obtain sufficient SNR to answer definitively this fundamental question.

A Pandora box of science

Given its efficiency, the high spectral resolution - and in particular the
R = 200’000 mode -, the extreme high radial-velocity precision, and the possibility of combining up to 4 UTs, ESPRESSO will not only be able to fulfill the main scientific objectives, but also open a new parameter space in observational astronomy with hopefully many new and unexpected results. The original scientific report of ESPRESSO (Phase A) describes in detail additional unique and exciting sciences cases which can be addressed with ESPRESSO by the astronomical community.

The Science

Instrument and Operations

60 s

1200 s

60 s

3600 s

1200 s

3600 s

The implementation of the feed from the 4 UTs indirectly provides another major advantage for the 1-UT mode: operational flexibility. In 1-UT mode, ESPRESSO can be fed by any of the four UTs, which significantly extends the scheduling possibilities for ESPRESSO programmes and optimizes the use of VLT time in general. Scheduling flexibility is a fundamental advantage for survey programmes like radial velocity searches for extrasolar planets or time-critical programmes like studies of transiting planets. The 1-UT mode itself will thus greatly benefit from the implementation of the 4-UT mode.

In the 1-UT mode a SNR of 10 per extracted pixel is obtained in 20 minutes on a V = 16.3 star, or a SNR of 540 on a V = 8.6 star. We have estimated that at R = 130’000 resolution this SNR value will lead to 10 cm s-1 radial-velocity precision for a non-rotating, late-G or K star. In the 4-UT mode, the SNR of 10 on a V = 19.9 star is achieved with a single 1-hour exposure.

UT1

UT4

UT3

UT2

ESPRESSO – A new-generation spectrograph

ESPRESSO is a fiber-fed, cross-dispersed, high-resolution, echelle spectrograph. The telescope light is fed to the instrument via a Coudé-Train optical system and fibers. ESPRESSO is located in the Combined-Coudé Laboratory (incoherent focus) where a front-end unit can combine the light from up to 4 Unit Telescopes (UT) of the VLT. The target and sky light enter the instrument simultaneously through two fibers, one for the target and one for the sky, which form together the ‘slit’ of the spectrograph.

Several optical ‘tricks’ are used to obtain high spectral resolution and efficiency despite the large size of the telescope and the 1 arcscec entrance aperture of the instrument.  At the spectrograph entrance the an anamorphic pupil slicing unit (APSU) shapes the beam in order to compress the beam in cross-dispersion direction but not in main-dispersion direction, where the full resolving power is required. In the latter direction the pupil is sliced and superimposed on the echelle grating to minimize its size. The rectangular white pupil is then re-imaged and compressed. Given the wide spectral range and the large number of spectral channels, two large 90 mm x 90 mm CCDs are required to record the full spectrum. Therefore, a beam splitter separates the beam in a blue and a red arm, which in turn allows to optimize each arm for image quality and optical efficiency. The cross-disperser has the function of separating the dispersed spectrum in all its spectral orders. In addition, an anamorphism is re-introduced to make the pupil square and to compress the order width, such that the inter-order space and the SNR per pixel are both maximized.

Feeding and operating the instrument with up to 4 UTs

ESPRESSO is located in the Combined Coudé Laboratory of the VLT (CCL). It is fed from any of the UT via the M4 mirror at Nasmyth (1-UT or singleHR mode). A fully optical system transports the beam through the telescope’s Coudé path to the Coudé focus and along the ‘incoherent’ tunnels to the Combined-Coudé Laboratory. This location opens the unique opportunity to feed it with the four UTs simultaneously, thus yielding a 16m-equivalent telescope. As such, the 4-UT mode (multiMR) will offer an important gain in collecting area and in read-out noise, since heavy binning can be applied. Given the operational constraints, it appeared interesting to offer a "reduced" 4-UT mode in which only 2 or 3 UTs are fed to ESPRESSO. Operationally the instrument will have only two main modes, a 1-UT and a 4-UT mode. In the 1-UT mode ESPRESSO will behave like any other VLT instrument. In the 4-UT mode, one can choose to feed any 2, 3 or 4UTs into the instrument. The spectral resolution will be reduced by a factor of two (given twice the collecting area diameter), but efficiency and field size will be conserved throughout the chosen configuration.

Extreme spectroscopic and radial-velocity precision

The extreme precision will be obtained by adopting and improving well-known HARPS concepts. The light of 1 or several UTs is fed by means of the front-end unit into optical fibers that scramble the light and provide excellent illumination stability to the spectrograph. In order to improve light scrambling non-circular fiber shapes will be considered. (Laboratory tests on octagonal fibers have demonstrated that the scrambling efficiency can be increased by a factor of 10.) The target fiber can be fed either with the light of the astronomical object or the calibration sources. The reference fiber will receive either sky light (faint source mode) or calibration light (bright source mode). In the latter case - the famous simultaneous-reference technique adopted in HARPS - it will be possible to track instrumental drifts down to the cm s-1 level. This technique relies on the perfect knowledge and stability of the spectral lines of a spectral reference source. Therefore, ESPRESSO will integrate a new developed laser frequency comb that provides stability and accuracy of better than 10-12, well below the mm s-2 regime.

| Parameter/Mode | singleHR (1 UT) | multiMR (up to 4 UTs) | singleUHR (1 UT) |
| --- | --- | --- | --- |
| Wavelength range | 380-780 nm | 380-780 nm | 380-780 nm |
| Resolving power | 130’000 | 55’000 | 200’000 |
| Aperture on sky | 1.0 arcsec | 4x1.0 arcsec | 0.5 arcsec |
| Spectral ampling (average) | 4.5 pixels | 5.5 pixels (binned x2) | 2.5 pixels |
| Spatial sampling per slice | 9.0 (4.5) pixels | 5.5 pixels (binned x4) | 5.0 pixels |
| Simultaneous reference | Yes (no sky) | Yes (no sky) | Yes (no sky) |
| Sky subtraction | Yes (no sim. ref.) | Yes (no sim. ref.) | Yes (no sim. ref.) |
| Total efficiency | 12% | 12% | TBD |
| Instrumental RV precision | < 10 cm s-1 | ~ 1 m s-1 | < 10 cm s-1 |

A 1 nm physical shift of the spectrum corresponds to 5 cm s-1 radial-velocity shift. Although the simultaneous reference is able to track potential drifts, the position of the target spectrum on the CCD shall be maintained as stable as possible in order to avoid higher-order effects. The whole spectrograph is therefore installed in a vacuum chamber and controlled in temperature. Pressure and temperature variation will be kept below 0.001 mbar and 0.001 K, respectively, such that the spectrum will not drift by more than 10 cm s-1 on daily basis. The detector system, which is based on a standard ESO model, will as well be optimized for thermal and mechanical stability.

End-to-end operation

The overall efficiency and the scientific output of long-lead programmes can be considerably increased if an integrated view of the operations is adopted. ESPRESSO shall not be considered as a stand-alone instrument but as a ‘science-generating machine’. An integrated view can only be obtained if the operations consider all the aspects, from the operation preparation through the instrument operation and control SW and HW, to the data reduction and analysis, and the related data products. ESPRESSO will deliver full-quality scientific data less than a minute after the end of an observation. These data, including the reduced data, will be archived and periodically reprocessed at ESO to obtain a coherent data set.

All these steps will be optimally integrated into the Paranal environment such that the result is satisfactory for the observatory and the Users. The final goal is to provide eventually the User with scientific data as complete and precise as possible.

Instrument and Operations

### Management

The final consumer of ESPRESSO is the astronomical community of the member states of ESO. The customer is ESO, which specifies the technical requirements for ESPRESSO.  The consortium is the contractor of the instrument. It is formed of institutes of four countries. Each country is represented in the Consortium by a single point of contact for all managerial aspects, the Co-PI. The Consortium is represented in the contractual relationship with the customer, ESO, by a single point of contact, the PI. The legal aspects within the Consortium are regulated by an ‘internal’ Memorandum of Understanding (MoU), which describes obligations and rights of the single Institutes within the project, and defines the use of and rights over the scientific outputs of ESPRESSO.

| Country | Institute | Responsible |
| --- | --- | --- |
| Switzerland | Observatoire de Genève, University of Geneva | Francesco Pepe, PI |
|  | Physikalisches Institut, University of Bern |  |
| Italy | INAF, Osservatorio Astronomico di Trieste | Stefano Cristiani, Co-PI |
|  | INAF, Osservatorio Astronomico di Brera |  |
| Spain | Istituto de Astrofìsica de Canarias, Tenerife (IAC) | Rafael Rebolo Lopez, Co-PI |
| Portugal | Centro de Astrofìsica, University of Porto (CAUP) | Nuno Santos, Co-PI |
|  | Faculdade da Ciências, University of Lisbon (FCUL) |  |

The relationship with ESO is somewhat more complex because ESO is, on the one hand the customer, and, on the other hand, project member. This Partnership is expressed by the fact that it holds the responsibility for various work packages and deliverables. In order to formalize the collaboration ESO has signed the mentioned MoU as an ‘Associated Partner’ with rights and obligations similar to the other Partners. ESO’s work packages are bound to a unique project structure and their responsible persons communicate directly with the ESPRESSO Project Manager. The ESO’s Representative is part of the Executive Board which is composed of the Consortium’s Co-PIs.  The Executive Board is responsible for all decisions which concern contractual matter and the management structure of the Consortium.

Project structure and phasing

The project is led by the project manager who defines the project structure (work-break down), assigns responsibilities, defines a schedule and is responsible for project controlling. His is supported by a System Engineering Team, which is responsible for the various system aspects and interfaces. Another team, the Science Advisory Team, will continuously verify that the project evolves in the direction of the declared scientific goals and that the instrument capabilities will be eventually compliant with them. Moreover, the Science Advisory Team provides support in scientific questions such as operation and calibration, observations and scheduling, requirements definition, etc.

Project and Management

Project and Management

Project structure and phasing

The project is broken down into 5 phases – Preliminary Design, Final Design, Subsystem

### Schedule

The official project kick-off was held in February 2011. The design phase will last about 2.5 years and is planned to end in Summer 2013. The procurement of components and manufacturing of subsystems will last about 18 months. In order to allow for early procurement of critical components a separate FDR for these long-lead items (LLI) will be held in advance. In the second half of 2014 the subsystems will be ready for integration in Europe at the integration site. The AIV phase will last about 12 months and end in the 3rd quarter of 2015. The four Coudé-Trains will be tested in Europe separately and brought, after acceptance, directly to the VLT. Verification of its performances and interfaces will be carried out before the spectrograph arrives at the VLT. Acceptance Europe is foreseen to be held in Fall 2015. The transfer of the instrument to Paranal, installation and on-site commissioning will take place in 2016. Acceptance Paranal is finally planned by the end of 2016.

### Project costs

The general concept was adopted that each of the Partners will carry capital and human costs associated to the work packages which they are responsible for. The institutes have committed to raise capital funds at a level of about 1.5 MEUR per country and to provide the needed manpower resources. The Consortium’s capital expenses including travel and overheads have been estimated at the level of 9 MEUR and the human effort to about 113 FTE. ESO’s contribution is 3.6 MEUR in capital and about 13 FTE in manpower.

| Milestone | From TP | Planned Date |
| --- | --- | --- |
| KOM |  | 4-5/10/2010 |
| TKO | TP=01/02/2011 | 17-18/01/2011 |
| PDR | TP + 10 months | 29-30/11/2011 |
| LLI FDR | TP + 17 months | 07/ 2012 |
| HW and SW FDR | TP + 26 months | 04/ 2013 |
| IRM | TP + 38 months | 04/ 2014 |
| IRRs |  |  |
| -SS tbd | TP + 38-50 | 04/2014 - 04/2015 |
| -grating+red arm | TP + 44 | 10/2014 |
| -blue arm | TP + 50 | 04/2015 |
| Coudé-Train PAE | TP + 52 months | 06/2015 |
| Front-End PAE | TP + 56 months | 10/2015 |
| Instrument PAE | TP + 56 months | 10/2015 |
| PAC | TP + 70 months | 12/2016 |
| FAC | TP + 94 months | 12/2018 |

|  | Phases A-C | Phases D-E | Total |
| --- | --- | --- | --- |
| Manpower | 7'310'221 € | 2'776179 € | 10'086'400 € |
| FTE | 91.38 | 34.70 | 126.08 |
| Travels | 954'971 € | 503'529 € | 1'458'500 € |
| Services | 79'857 € | 197'143 € | 277'000 € |
| Overheads | 1'907'568 € | 797'232 € | 2'704'800 € |
| Capital costs | 8'252'223 € | 111'877 € | 8'364'100 € |
| Total | 18'504'841 € | 4'385'959 € | 22'890'800 € |

The project is broken down into 5 phases – Preliminary Design, Final Design, Subsystem MAIV, System AIV, and AIV and Commissioning in Paranal. Each phase contains several work-packages, which describes the sum of all activities necessary to successfully complete the project. Each work package is assigned to one institute and one single responsible person only. The Co-PI responsible for this institute commits to provide the WP manager with all resources necessary to complete the work package successfully.

ESO

### Appendices

Appendix A: 	List of PDR Documents

Appendix B: 	List of Applicable Documents

Appendix C: 	List of Reference Documents

Appendix D: 	List of Acronyms

Appendix E: 	List of Names and Abbreviations

ESO

October 2011

The ESPRESSO Consortium

ESPRESSO:

A super-HARPS on a 10 m-class telescope

A spectral coverage from 380 to 780 nm in one shot

The highest-resolution instrument on a 10 m-class telescope

A wavelength calibration far more accurate than any other facility

An instrument producing cleanest, best-quality spectra, both at high and low SNR

A spectrograph on a 16 m telescope, the largest visible photon-collector until ELTs will be available

An ultra-high resolution mode (R~200,000), far beyond other existing facilities on a 10 m-class telescope